## Sheet1
Defibrillator Analyzer
1 | DESCRIPTION
1.1 | An analyzer is used to test the output performance of the defibrillator.
2 | TECHNICAL SPECIFICATION
2.1 | The unit offered must be an auto ranging Defibrillator Analyser which must include a built in 12-Lead ECG Simulator and also a built in External Non-Invasive Pacer Tester capable of displaying test verification data.
2.2 | The Analyzer offered must be as versatile as possible, portable and must be suitable for testing and verifying all makes and models of Defibrillators, including various AED’s, and must be supplied with a carry case.
2.3 | The Analyzer offered must be easy to use (user friendly) with a menu key interface allowing quick access to all tests
2.4 | The Analyzer must have a reasonable size display for display of test values in the English Language and this display must be clearly visible under all lighting conditions.
2.5 | The Analyzer must be capable of measuring delivered energy from all different technologies of Defibrillators (Monophasic, biphasic and pulsed-biphasic defibrillation technology), in watt seconds (joules) with a simulated resistance of typical 50 Ohm impedance
2.6 | The Analyzer must be able to measure Sync time
2.7 | The Analyzer must be able to generate waveforms as follows:
2.7.1 | Normal sinus rhythm 30, 60, 80, 120, 200 and 250 beats per minute
2.7.2 | Sinewave 10, 40, 50, 60 and 100 hertz
2.7.3 | Arrhythmias – Atrial Fibrillation
2.7.4 | Arrhythmias – Atrial Flutter
2.7.5 | Arrhythmias – Atrial Tachycardia
2.7.6 | Arrhythmias – Ventricular Fibrillation (Coarse and Fine)
2.7.7 | Arrhythmias – Supraventricular Tachycardia
2.7.8 | Arrhythmias – Atrial PAC
2.7.9 | Arrhythmias – Multifocal PVC’s
2.7.10 | Arrhythmias – Asystole
2.8 | The Analyzer offered must deliver waveform outputs via 12-lead ECG integrated posts and also paddle plate adapters
2.9 | The Analyzers built in pacer tester, with preferable pacer brand selection, must measure and display, pulse rate, pulse width, current, energy and refractory periods with a selectable input load of typical 50 Ohm to 1000 Ohm
2.1 | The unit offered must have an accuracy of + 1.5% at auto range reading of 1 – 600 joules
2.11 | The Analyser must be able to make output power measurements at:
2.11.1 | Voltage Range: 50V to 5000V
2.12 | The Analyser must be able to make output power measurements at:
2.12.1 | Current Range: 1A to 100A
2.13 | The unit must have a communication port / connectivity (Wired or Wireless), which will enable connection to a PC for automation test software and firmware/waveform upgrades
2.13.1 | Bidder to provide the name/make of Automation Test Software as stated above at 2.13
2.14 | The automation test software, used with the unit, must be able to integrate with various Computerised Maintenance Management System (CMMS) software’s, which is an advantage for future implementation
2.14.1 | Bidder to provide the name and make of various CMMS programs currently compatible with the unit.
2.15 | The Analyzer must be able to operate on 220 Volt + 10% 50hz single phase a.c. power supply. The Analyzer must be surge protected.
2.16 | The Analyzer offered must have an optional external Test Load that provides multiple impedance loads of 25 Ω, 50 Ω, 75 Ω, 100 Ω, 125 Ω, 150 Ω, 175 Ω, and 200 Ω to comply with IEC 60601-2-4 Standard. The description, price and catalogue number must be clearly stated on the schedule
2.17 | Bidder to indicate which consumables and non-consumable items are required for the normal operation and standard maintenance of the offered Analyser and tenderer to specify the cost of these items on the attached schedule should this clause be applicable
2.18 | Bidder to specify what the approximate present calibration cost on the offered Analyzer will be, and the frequency/service interval of the calibration required for the tendered Analyser
2.19 | All electrical components of the Analyser must be modular and its design shall be such that liquid splashes will not gain entry into the P.C. Boards
2.2 | The Analyzer offered must be supplied with a complete set of accessories as part of the unit, at no extra cost, in order that the Analyser may be put to immediate use. All those accessories that will be supplied as part and parcel of the Analyzer must be specified. The unit must come in a carry case.
2.21 | Must be provided with a manual guide preferably in the English language.