## Sheet1
Electrosurgical Analyzer
1 | DESCRIPTION
1.1 | Used to test perfomance of diathermy unit.
2 | TECHNICAL SPECIFICATION
2.1 | The unit offered must have the latest up to date technology, it must be fully portable and it must be supplied with a soft carry case
2.2 | The unit offered must be user friendly and offer greatest ease and safety in operation and be able to test both Monopolar and Bipolar modes of all the various makes and models of Electrosurgical Machines
2.3 | The unit offered must be able to provide function test verification to all makes and models of Electrosurgical Machines according to Manufacturers original Specifications and also to I.E.C. 601-2-2 Standard
2.4 | The unit offered must be available with internal operator selectable test load variables in a minimum range of 10 ohms to 5200 ohms and must be adjustable for both Monopolar and Bipolar modes
2.5 | The unit offered must be capable to run an Automatic Power Distribution measurement test, which includes power, current, peak-to-peak voltage, and crest factor values through a user selectable load range
2.6 | The unit must be able to perform High-frequency leakage measurements with accuracy of +/- 2%
2.7 | The Analyzer must have a reasonable size display for display of test values in the English Language and this display must be clearly visible under all lighting conditions.
2.8 | The unit must be able to carry out power measurements in a minimum range of 1 Watt up to 500 Watts and this power measurement must be autoranging
2.9 | The unit must be able to perform current measurements in a minimum range of 30 to 2200 milliamps RMS with an accuracy of +/- 2% of readings
2.1 | The unit must be provided with an isolated 200 Ohm fixed load for I.E.C. testing
2.11 | The unit offered must be able to perform HF current leakage test, in a measured range of 30 to 2200mA, with 200 Ohm fixed load from patient plate, active electrode and chassis, as per IEC 601-2-2
2.12 | The unit’s frequency response must be in a typical range of 30Hz to at least 5MHz.
2.13 | The unit must be fully functional on 220 Volt 50 Hz single phase a.c. power supply and must be fused in both the Live and Neutral and the fuses must be accessible without the need to strip the unit
2.14 | Unit offered must be supplied with an isolated BNC connector for coupling to an oscilloscope
2.15 | The unit tendered for must have a built in power supply
2.16 | The tenderer must ensure that all leads; cables and accessories required for testing must be supplied as standard, at no extra cost, with the Analyzer so that it could be put into immediate use
2.17 | The unit must have a communication port / connectivity (Wired or Wireless), which will enable connection to a PC for automation test software and firmware upgrades
2.18 | Bidder to state the name/make of Automation Test Software. The price must be added to the schedule of optional accessories
2.19 | The unit must have Automation Test Software. The bidder must state the name/make of Automation Software.
2.2 | The automation test software, used with the unit, must be able to integrate with various Computerised Maintenance Management System (CMMS) software’s, which is an advantage for future implementation
2.21 | Bidder to provide the name and make of various CMMS programs currently compatible with the unit
2.22 | The bidder must undertake to provide the Health Technology Services – Clinical Engineering’s in house Technicians training on the use of the tendered product in order that the Technicians become fully acquainted with the Technical support capability of the Analyzer
2.23 | The Analyzer must be able to operate on 220 Volt + 10% 50hz single phase a.c. power supply. The Analyzer must be surge protected.
2.24 | The Analyzer offered must be supplied with a complete set of accessories as part of the unit, at no extra cost, in order that the Analyzer may be put to immediate use. All those accessories that will be supplied as part and parcel of the Analyzer must be specified. The unit must come in a carry case.
2.25 | Must be provided with a manual guide preferably in the English language.