## Sheet1
Infusion Pump Analyzer
1 | DESCRIPTION
1.1 | Infusion pump analyzers are test devices that aid in determining the performance of infusion pumps.
2 | TECHNICAL SPECIFICATION
2.1 | The Analyzer offered must be microprocessor controlled and must be able to test all makes and types of infusion devices presently in use at all institutions, which must include syringe driver pumps and volumetric infusion pumps
2.2 | The Analyzer offered must be supplied as a unit with multiple channels operation but must be upgradeable to multichannel if required in the future so that more than two infusion devices could be tested simultaneously
2.3 | If the Analyzer is ordered as a multichannel unit, then each channel must work independently of each other
2.4 | The Analyzer must be capable of measuring both steady and non-steady flow patterns and be capable of displaying Flow and Volume measurements
2.5 | The Analyzer must have a reasonable size display for the display of test values in the English Language and this display must be clearly visible under all lighting conditions.
2.6 | The Analyzer must be capable of carrying out occlusion pressure measurements of 0 to at least 2500 mmHg.
2.7 | The unit must have a communication port / connectivity (Wired or Wireless), which will enable connection to a PC for automation test software and firmware upgrades
2.8 | The automation test software, used with the unit, must be able to integrate with various Computerised Maintenance Management System (CMMS) software’s, which is an advantage for future implementation
2.9 | Bidder to provide the name and make of various CMMS programs currently compatible with the unit
2.1 | The Analyzer must have on board memory for storing of test results and for printing or downloading to a computer
2.11 | The Analyzer must be capable of carrying out Back Pressure measurements in a typical range of -100mmHg to + 300mmHg
2.12 | All electrical components of the Analyzer must be of modular design and the Analyser must be designed such that liquid splashes will not gain entry into the P.C. Boards
2.13 | The average flow rate measurement range of the Analyzer must be in the typical range of 0.1 to1500ml/h with an accuracy of + 1% at flow rates of between 20 and 200ml/h and + 2% at all other flow rates
2.14 | The volume delivery measurements range of the Analyszer must be in the typical range of 1 to 9999 ml accuracy of + 1% for a delivered volume of greater than 20ml and + 2% for a delivered volume of less than 20ml
2.15 | The PCA measurement must be available as follows(The accuracy must be + 2% of reading and + 1 LSD after 10ml delivered volume):
2.15.1 | Minimum bolus volume of at least 0.5ml
2.15.2 | Maximum test duration of at least 100 Hours
2.15.3 | Maximum bolus volume of at least 9999ml
2.16 | The time taken by the Analyser before producing the first flow rate measurement must be typically as follows:
2.16.1 | 1ml/hr + 5 minutes
2.16.2 | 100ml/hr + 25 seconds
2.16.3 | 1000ml/hr + 3 seconds
2.17 | The Analyzer must be able to operate on 220 Volt + 10% 50hz single phase a.c. power supply. The Analyzer must be surge protected.
2.18 | The Analyzer offered must be supplied with a complete set of accessories as part of the unit, at no extra cost, in order that the Analyzer may be put to immediate use. All those accessories that will be supplied as part and parcel of the Analyzer must be specified. The unit must come in a carry case.
2.2 | The frequency of calibration required for the Analyzer and the present cost of calibration must be specified
2.21 | Must be provided with a manual guide preferably in the English language.