## Sheet1
MultiGas Analyzer
1 | DESCRIPTION
1.1 | A gas analyzer is an analytical instrument used to identify the concentration of known gases in an atmosphere that contains multiple gases
2 | TECHNICAL SPECIFICATION
2.1 | The Analyzer must be specifically designed for highly accurate measurement and verification of the quality of piped N2O and O2 / (Medical Air) gases in hospitals, as well as identification of contaminants CO and CO2, and must be supplied with a convenient hard carry case
2.2 | The Analyzer offered must be as versatile as possible, portable, lightweight and have a protective carrying case.
2.3 | The Analyzer offered must be user friendly, with a menu and numerical keypad interface, allowing quick access to all measurements and tests, and must display all measured parameters simultaneously on the main display
2.4 | The Analyzer offered must be able to connect to any medical gas outlet (O2, Air and N2O) in the hospital to quickly verify the concentration and quality of the supplied gas
2.5 | Bidder to supply necessary accessories and regulator kit to connect to various gas outlets and must be supplied as part of the final offer at no extra cost
2.6 | The Analyzer offered must be capable of being calibrated by the end user with relative ease. Calibration gas canister must be provided with the unit.
2.7 | The Analyzer offered must be able to store any current reading/measurement, for ease of data capturing, and allow the data to be stored with an allocated ID, to identify the reading at a later point
2.8 | The Analyzer must be able to measure Oxygen and Air as follows:
2.8.1 | Measurement Range: 0% to 100% (Air = 21% of O2 reading)
2.8.2 | Accuracy: +/- 2% full scale
2.9 | The Analyzer must be able to measure N2O as follows:
2.9.1 | Measurement Range: 0% to 100%
2.9.2 | Accuracy: +/- 2%
2.1 | The Analyzer must be able to measure CO2 as follows:
2.10.1 | Measurement Range: 0 to 2000ppm
2.10.2 | Accuracy: +/- 2%
2.11 | The Analyzer must be able to measure CO as follows:
2.11.1 | Measurement Range: 0 to 500ppm
2.11.2 | Accuracy: +/- 5%
2.12 | The Analyzer must be able to measure the gasses, and produce a stable result, within the following response times:
2.12.1 | N2O: No longer than 40 seconds
2.12.2 | O2 / Air: No longer than 60 seconds
2.12.3 | CO2: No longer than 40 Seconds
2.13 | Bidder must state the scientific method used by the analyser to measure the different gasses
2.14 | Bidder to specify if any replacements of any measurement cells are required, and at what interval, and the cost thereof must be clearly indicated on the schedule of optional accessories
2.15 | The Analyzer must be able to operate on 220 Volt + 10% 50hz single phase a.c. power supply. The Analyzer must be surge protected.
2.16 | The Analyzer offered must be supplied with a complete set of accessories as part of the unit, at no extra cost, in order that the Analyser may be put to immediate use. All those accessories that will be supplied as part and parcel of the Analyzer must be specified. The unit must come in a carry case.
2.17 | Must be provided with a manual guide preferably in the English language.