## Sheet1
Ultrasound Watt Meters
1 | DESCRIPTION
1.1 | An Ultrasound Watt-meter (Radiometer) is a testing device used in performing output measurement tests and performance verification on ultrasound equipment.
2 | TECHNICAL SPECIFICATION
2.1 | The unit must be portable, it must be a precision wattmeter and employ a simple and accurate technique for carrying out power output measurements for therapeutic ultrasound devices (physiotherapy). Please provide details on the measurement technique employed by this unit
2.2 | The unit must be able to measure the output of all makes and models of therapeutic ultrasound (Physiotherapy) devices in a minimum frequency range of 1 to 10MHz.
2.3 | The unit must have a minimum power output measurement range of 0 to at least 30W.
2.4 | The unit must have a high resolution of at least 10mW.
2.5 | The unit must be provided with a mechanism that will ensure and hold the Therapeutic transducer in place whilst performing test power output measurements.
2.6 | The unit must be equipped with HDMI and/or USB computer interface ports.
2.7 | The unit must have a software integration that allows for the download and printing of results.
2.8 | The unit must be able to carry out both PULSED and CONTINUOUS power output measurements of therapeutic ultrasound units (physiotherapy).
2.9 | A digital readout to display the measurements in watts must be provided .
2.9.1 | The Meter must have a reasonable size display for display of test values in the English Language and this display must be clearly visible under all lighting conditions.
2.9.2 | The DISPLAY should be LCD or LED type.
2.1 | Where the unit is able to operate off the 220 Volt 50hz a.c. power, it should be clear whether the unit works off an a.c. adapter and also if it is supplied as a standard accessory.
2.10.1 | In the event where the a.c. adaptor is an optional extra, it must be supplied as part and parcel of the unit.
2.10.2 | It should be stated if the unit is able to function off battery power. The voltage and type of battery and whether it is re-chargeable, should be stated.
2.11 | The accuracy of the unit should be known.
2.12 | The unit must be able to accommodate a range of transducers with various diameters. The various diameters of the transducers that could be accommodated for power output measurements must be known.
2.13 | The unit must have a display zeroing facility state whether this zeroing is automatically done or whether it requires operator intervention.
2.14 | The unit must be supplied with a carry case. It must be stated whether the carry case is a SOFT or HARD type.
2.15 | The unit must be well protected against liquid splashes gaining entry on to the internal P.C. Board etc.
2.16 | The Meter offered must be supplied with a complete set of accessories as part of the unit, at no extra cost, in order that the Meter may be put to immediate use. All those accessories that will be supplied as part and parcel of the Meter must be specified. The unit must come in a carry case.
2.17 | Must be provided with a manual guide preferably in the English language.