## Sheet1
Ventilator Simulator
1 | DESCRIPTION
1.1 | Ventilator simulator test gas flow and respiratory medical equipment accurately
2 | TECHNICAL SPECIFICATION
2.1 | The unit must be capable of testing all models of ventilators virtually: adult, paediatric, anaesthesia,neonatal and high-frequency.
2.2 | The unit must also test a variety of other medical equipment such as anaesthesia gas delivery machines, insufflators and oxygen concentrators.
2.3 | The unit must be able to operate on both:
2.3.1 | batteries -minimum of 8 hrs continuously
2.3.2 | Mains - the simulator must be able to operate on 220 Volt + 10% 50hz single phase a.c. power supply. The simulator must be surge protected. The unit mains cable must be fitted with a SABS approved 15 amp plug top.
2.4 | The simulator must have a reasonable size display for display of test values in the English Language and this display must be clearly visible under all lighting conditions.
2.5 | Single full range channel
2.6 | Ultra low flow port :+/- 750ml/min
2.7 | Ultra low pressure port: 0 to 10 mbar
2.8 | Flow: a)range :+/- 300 slpm b)air accuracy: 1.7%
2.9 | Barometric pressure range of 550 to 1240 mbar @ +/-1%
2.1 | Temperature: 0 to 50 degrees
2.11 | Humidity: 0 to 100% @ +/-3%
2.12 | Oxygen: 0 to 100% @ +/-1%
2.13 | Minimum Gas parameters for testing :
2.13.1 | Oxygen
2.13.2 | Medical air
2.13.3 | N2O
2.13.4 | CO2
2.14 | Inspiratory tidal volume :0 to 60l
2.15 | Expiratory tidal volume range :0 to 60l
2.16 | Minute volume range 0 to 100l
2.17 | Breath rate range : 1 to 1500 bpm
2.18 | Inspiratory to expiratory time ratio (I:E) range: 1:300 to 300:1
2.19 | Peak inspiratory pressure range : +/-160 mbar
2.2 | Positive end inspiratory pressure range : +/-160 mbar
2.21 | Lung compliance range :0 - 1000ml/mbar
2.22 | inspiratory time range: 0 to 60 s
2.23 | Espiratory time range: 0 to 90 s
2.24 | Peak expiratory flow range : +/-300 lpm
2.25 | Peak inspiratory flow range : +/-300 lpm
2.26 | The simulator offered must be supplied with a complete set of accessories as part of the unit, at no extra cost, in order that the simulator may be put to immediate use. All those accessories that will be supplied as part and parcel of the simulator must be specified. The unit must come in a carry case.
Ventilator Simulator - Accessories
1 | The unit must be supplied with at least 1L test lung
2 | The lung must present the following restrictors:
2.1 | Rp5 - PRESSURE DROP 10.80 cmH20
2.2 | Rp20 - PRESSURE DROP 17.60 cmH20
2.3 | RP50 - PRESSURE DROP 27.00 cmH20