## Sheet1
320 Slices CT Simulator
Item | Technical Specifications | Weight
System should be compatible with or allow for inclusion of surface guided and/or image guided/adaptive radiotherapy systems.
1 | The system must be a large Bore CTSim for radiotherapy use (at least 80 cm)
1.1 | The system must have 320 Slices, producing 320 Slices per 360 degrees rotation
1.2 | The unit must be low voltage slip-ring designed gantry technology.
1.3 | The unit must support whole body CT Scanning.
1.4 | The unit must have a minimum scan time of 0.4 second or less for 360 degrees rotation
1.5 | The unit must have a gantry aperture size of at least 800mm or larger.
1.6 | Angulations must at least + 30degrees and - 30 degrees.
1.7 | The Unit must be able to scan either head first or feet first.
1.8 | The unit must have a minimum field of view at least 500mm or larger.
1.9 | The unit must have control panels on either side (left and right) of table on the gantry to adjust gantry movements.
1.1 | The system must have triple laser lights for accurate patient positioning at setup plane. (Accuracy of 1mm)
1.11 | Control of gantry movements from operator console must be possible.
1.12 | The system shall have an Intercom between gantry and operator console.
1.13 | Electronic patient breathing instructions in multiple languages must be possible.
1.14 | Emergency stop switches to be available on gantry as well as on the console.
2 | PATIENT SUPPORT SYSTEM (PATIENT TABLE)
2.1 | Support patients 220kg or more.
2.2 | Lowest table height must be 550mm or less
2.3 | The table length must be 2 meters or more
2.4 | Longitudinal table top travel must not be less than 1700mm for whole body coverage.
2.5 | Must have motorized patient support movements (longitudinal and vertical movements)
2.6 | A metal-free table top over scanner range allowing patient scanning without re-positioning patient must be provided.
2.7 | Control of table movements from operator console and gantry must be available.
2.8 | Motorized removal and re-indexing of table top without repositioning patient scanning for emergency access to patient must be available.
2.9 | Patient positioning accessories required to perform the procedures specified, including headrest support, axial and coronal head rest support, table leg extender, security straps, infant immobiliser, arm support, knee support and immobilizing straps must be provided
2.1 | Emergency override to enable manual removal of patients in case of power failure or emergency must be included
3 | DETECTOR ACQUISITION SYSTEM
3.1 | Solid-state detector technology using low-dose and high resolution acquisitions.
3.2 | Multi-slice detector technology for 64-slice acquisitions per 360 degree gantry rotation.
3.3 | Multi-Slice detector with 1200 views per second sampling rates or better.
3.4 | Pre-body collimator must be available.
3.5 | The ability to select 64-slice acquisition with 0.6mm Slice thickness or thinner. State number of channels
3.6 | All calibration and service phantoms with wall brackets must be included
4 | X-RAY GENERATOR
4.1 | Modern high frequency generator allowing at least between 80kV to 135kV and 500mA or better.
4.2 | To support at least 200 scans in 100 seconds using a 0.5 second scan time or better during the spiral scan mode.
4.3 | The unit must be able to scan continuously a length of 1.7m high resolution without cooling time.
4.4 | The unit must support repeat scanning with minimal wait time.
4.5 | The unit must have at least 3 variable kV setting between 80kV (or less) and 135kV (or more).
4.6 | The unit must have variable mA settings up to at least 500mA or better.
4.7 | State maximum mA allowed for each kV.
4.8 | Generator Power output of at least 60kW or higher. State actual power output.
5 | X-RAY TUBE
5.1 | The unit must have a high performance rotating anode tube with dual focal spot.
5.2 | The unit must have a total Filtration of 3.5mm AI equivalent or greater.
5.3 | Anode heat dissipation rate not less than 1300kHU/minute.
5.4 | Cooling performance curves required.
5.5 | The unit must have a heat storage capacity of at least 7MHU.
5.6 | Expected tube life. State in years and slices
6 | OPERATOR CONSOLE
6.1 | The operator console must have a display and entry of patient data pre-planning
6.2 | The console must have comprehensive adult and paediatric anatomical programs/protocols.
6.3 | Must at least have 100 pre-programmed scan protocols required with option to add user defined protocols.
6.4 | One operator console and two independent post -processing workstations must be included in the bid.
6.5 | The operator console to include a high speed computer which will support multi-tasking operation to allow simultaneous scanning reconstruction, viewing, archiving, processing and filming operation.
6.6 | Image review system must be able to display, analyse, process and manage acquired imaging data such as:
6.6.1 | - Zoom
6.6.2 | - Grey scale (colour display) adjustment
6.6.3 | - Measurement of distance and angle
6.6.4 | - Rotation
6.6.5 | - Grid
6.6.6 | - Annotation
6.6.7 | - Cine display
6.6.8 | Facilities must be available to perform basic CT examination of all parts of the body, including reconstruction, review and printing.
6.7 | All post processing software must be available on all three workstations
6.8 | Compatibility: To comply to DICOM 3 standards. Please provide DICOM 3 conformance statement.
6.9 | On line storage: - To store at least 800 Tera Bytes uncompressed images of 512 x 512 images.
6.1 | Archival storage by means of an DVD recording system must be provided
6.11 | X-Ray tube heat loading display is required on console
6.12 | Print facility to laser imager must be supported from all CT workstations.
3 | Chairs: wheeled, height adjustable, swivel for the operators must be supplied. (Quantity 4)
6.14 | All high end post processing software must be on both the CT console and Doctors console. (The user must be able to either configure the post processing software on CT console or Doctors console depending on workflow requirements.)
7 | COMPUTER SYSTEM / RECONSTRUCTION
7.1 | At least a 512 x 512 reconstruction matrix is required.
Details of matrix sizes available to be stated
7.2 | Real time reconstruction for scout-view required.
7.3 | A 0.1 second or faster reconstruction time is required per 512 x 512 reconstructed image
7.4 | 3 - D re-construction must be possible
8 | IMAGE STORAGE/Export
8.1 | Large image archive to store raw scan-data of at least 1000 GB as well as reconstructed data required. State size
8.2 | At least 3600 scan data sets are to be stored. State storage size.
8.3 | The Unit must be able to burn images onto CD & DVD. (200 CDs must be included)
8.4 | Dicom Images must be exported to Radiotherapy treatment planning system (¦Supplier must setup the export)
9 | IMAGE VIEWING MODES
9.1 | Comprehensive image acquisition, viewing, processing and filming modes required must be provided and supported by the systems.
9.2 | Noise reduction techniques must be available.
10 | SCAN MODES
10.1 | Scout-view mode for scan planning must be provided.
10.2 | Single scan mode must be provided.
10.3 | Scanning with serial reconstruction must be provided.
10.4 | Scanning with parallel reconstruction must be provided.
10.5 | Dynamic scan mode must be provided.
11 | SPIRAL MODE PERFORMANCE
11.1 | Tilted spiral scan mode with 64 slice and other combination for thin 0,6mm slice acquisition being acquired per gantry rotation must be provided.
11.2 | 10 scan plans per scout view required.
11.3 | The Unit must provide a variable pitch setting for spiral scanning mode.
11.4 | Single spiral sequence must be provided.
11.5 | Multiple spiral sequence must be provided.
11.6 | Combined conventional and spiral sequence must be provided.
11.7 | Tilted spiral mode must be possible.
11.8 | Real time reconstruction and display of images during the spiral mode must be provided at a viewing rate of at 16 images per second or better.
11.9 | For accurate contrast examinations, a bolus-timing / tracking package must be included.
11.1 | Isotropic volume imaging is required. State voxel sizes used for isotropic volume imaging.
11.11 | Standard mode spatial resolution of 15 lp/cm at cut-off or better is required.
State high contrast spatial resolution in line pairs per cm.
11.12 | High-resolution mode spatial resolution of 21 lp/cm at cut-off or better is required.
State high contrast spatial resolution in line pairs per cm.
11.13 | 4.0mm at 0.3% low-contrast resolution must be possible at 90mAs, 10mm slice thickness, 120kV.
State resolution and scan technique used to obtain measurement. "
11.15 | International CATPHAN performance phantom must be provided for measurements. Supply details used for measurements.
12 | DOSE REDUCTION TECHNIQUE
12.1 | Current dose reduction techniques used during patient scanning to reduce the total applied patient dose must be included.
12.2 | The applied patient dose must be quantified and displayed on the operators console for the patient examination.
12.3 | Bidders must include current dose reduction technology specifically to the following:
12.3.1 | - Gantry design.
12.3.2 | - Choice of scan-times available.
12.3.3 | - Detector design
12.3.4 | - Use of 64 x 0.6mm (or smaller) imaging techniques.
12.3.5 | - Slice profile used
12.3.6 | - Real-time exposure control during scanning
12.3.7 | - Variable pitch settings
12.3.8 | - Beam collimation and X-ray tube filtration
12.3.9 | - Adult and paediatric scanning exposure techniques/protocols.
12.3.10 | - Carbon fibre patient support table top attenuation
12.3.11 | - Accurate contrast timing.
12.3.12 | - Real-time spiral scanning display
12.3.13 | - Calculated dose display on operators console monitor.
12.3.14 | Software image quality correction that allows reconstructed images to be improved without having to rescan the patient with additional scans or scans with a higher applied dose.
12.3.15 | - For shoulders and pelvis with streak reduction techniques
12.3.16 | - Posterior fossa with image fusion techniques.
12.3.17 | - Abdomen with motion detection techniques.
12.4 | Any other techniques must be detailed.
12.5 | Supporting documentation indicating applied dose reduction must be included.
13 | DICOM 3 COMPATIBILITY
13.1 | The CT Scanner must support image & data linkage via network including all additional workstations and to laser camera. All DICOM software must be included in the cost. It must have the ability to be easily linked to any PACS and RIS system. It must have ability to be linked to ADSL/ISDN line or any other tele-radiology system
13.2 | All hardware & software relating to DICOM compatibility, including licences, must be supplied.
13.3 | The Laser camera to support DICOM compatible printing on 35 x 43 cm Laser film.
13.4 | Links to additional workstation and Laser camera systems is required through DICOM compatibility.
13.5 | DICOM Conformance statements on the various service classes per modality to be provided.
14 | THREE INDEPENDENT WORKSTATIONS ( One for Doctors reporting and two for operators control rooms)
14.1 | A WINDOWS Operating computer system is required with at least 4 GB RAM,
800 GB hard drive, network card, suitable graphics card and at least a 48cm flat panel LCD monitor (supporting at least 3 Megapixel resolution) is required.
14.2 | Provide modern user-interface with high speed-processing and storage facilities as per operators console.
14.3 | 2D, 3D, surface rendering, volume rendering, image sculpting, Maximum Intensity Projections, Minimum Intensity Projections, MPR and curved MPR and other image processing tools must be supported by system.
14.4 | All post processing software must be available on all workstations.
14.5 | Image manipulation and arithmetic must be available
14.6 | Archival storage using TIFF, JPEG, DICOM or similar format by means of a DVD must be included. (200 discs must be included)."
14.7 | Print facility to laser imager must be supported.
15 | DICOM COMPATIBLE LASER CAMERA
15.1 | DICOM Compatible networked laser camera must be provided.( 455 DPI or better, for printing on 35 x 43cm film size is required).
15.2 | 35 x 43cm laser film to be used
15.3 | Daylight loading film magazines to be used
15.4 | Linkage via network to additional workstation & CT Console must be provided.
15.5 | At least 100 films per hour printing to be possible
15.6 | Laser printing using dry printing technology required. Details of the unit to be offered.
15.7 | 35 x 43 cm film to be included in the bid price and delivered with the unit (500 sheets)
15.8 | UPS support must be included with at least 20 minutes back – up.
15.9 | All necessary hardware, software and licences must be included.
16 | SOFTWARE PACKAGES AND CLINICAL APPLICATIONS
16.1 | All packages required to perform the clinical applications as listed in SECTIONS 10.11 and 12. must be included.
16.2 | All inclusive Vascular, Oncology and orthopaedic clinical applications must be standard and included in the package. All other clinical applications as per list below must be available, however these must be quoted has optional:
16.3.1 | Virtual colonoscopy package with CO2 insufflator.
16.3.2 | CT Fluoroscopy for drainages, FNA biopsies, etc. is required, including the ability to interactively change table height, longitudinal table position and gantry angulation on remote hand control with ceiling mounted LCD monitor.
16.3.3 | Brain perfusion package.
16.3.4 | Cardiac package including patient monitor, ECG gated scanning, ECG gated reconstruction and full ventricle and coronary evaluation software.
16.3.5 | Neuro DSA software with total bone removal software.
16.4 | Any clinical application not listed in 6.2. but is offered by the manufacturer, must be quoted separately as an option.
17 | CT CONTRAST INJECTOR
17.1 | A modern dual head microprocessor controlled contrast injector is required for use with disposable syringes and will be used for all contrast enhanced examinations including cardiac and bolus chasing. (Please supply 200 units of each of the Consumables i.e. Syringes, Connecting Tubes, Y- Adaptor Tube with Dual check valve)
17.2 | The consumables for the injector mentioned above must readily available in South Africa
18 | POWER SUPPLY
18.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
18.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
18.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
18.4 | A 3 phase in / out power supply must be offered. State details.
18.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
18.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
18.7 | The lifetime of UPS must be at least 15 years.
18.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
18.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
18.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
18.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
18.12 | Maintenance free batteries must be offered.
20 | SUPPORTING DOCUMENTATION
20.1 | Sales licence issued in terms of Hazardous Substance Act must be submitted.
20.2 | All supporting documentation as requested in the various sections to be included.
20.3 | All original system brochures, product specifications and application notes to be supplied.
20.4 | DICOM 3 conformance statements to be included.
21 | CUSTOMER SUPPORT DETAILS
21.1 | The application training for radiologists and radiographers by a fully trained application specialist to include at least the following:
21.2 | CV of trainer to be supplied with the bid.
21.3 | An initial period of at least one week to be included patient scanning of multiple clinical problems.
21.4 | At least 4X 1week follow up periods during the first 12 months.
21.5 | Training in the specialized packages to be provided (in addition to the above training) Training schedule to be included. (to be discussed with End User)"
21 | STANDARDS AND SAFETY
21.1 | Should be FDA or CE approved product
21.2 | Electrical safety conforms to standards for electrical safety IEC-60601 / IS-13450
21.3 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
21.4 | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
21.5 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay)
21.7 | Lead Glass window to be installed as per room requirement
21.8 | Safety aspects of Radiation dosage leakage should be spelt out