## Sheet1
HIGH DOSE RATE (HDR) BRACHYTHERAPY UNIT
Item | TECHNICAL SPECIFICATIONS | Weight
1 | GENERAL DESCRIPTION
1.1 | High dose rate remote afterloader brachytherapy unit with radioactive source
2 | EQUIPMENT DETAILED SPECIFICATIONS
2.1 | Compulsory Technical specifications
2.1.1 | Remote after loading unit
2.1.2 | Iridium (Ir-192) preferred, Cobalt (Co-60), or equivalent high dose rate source
2.1.3 | Successful testing according to ISO 2919:2012-08
2.1.4 | Two dimension and three dimension (CT/MRI based ) treatment planning hardware, software's and license - latest in the market with most recent software version at the time of installation.
2.1.5 | Emergency source container
2.1.6 | Source replacement if Ir=192, to be included on warranty period.
2.2 | Compulsory Functional specifications
2.2.1 | Capable of delivering high dose rate remote afterload treatment.
2.2.2 | Two dimensional and three dimensional treatment planning unit.
2.2.3 | CT/MRI based and Dicom image compatible planning.
2.2.4 | Must interface with all Record & Verify systems for Oncology Management System (OMS) .
2.2.5 | Server based image import/export system which allows connectivity between CT/MRI simulator and HDR planning unit.
2.2.6 | In vivo dosimetry for real time patient monitoring (bladder and rectal doses).
2.3 | Compulsory Treatment delivery accessories : All applicators must be titanium (priced separetly as an option)
2.3.1 | Transfer tubes compatible with user applicators x18.
2.3.2 | Fletcher type, titanium CT/MRI shielded applicator set including intrauterine (including angulation variations) and ovoid set with varying cap diameters per set. X6
2.3.3 | Rectal retractor x6 (compatible with 2.3.2)
2.3.4 | Intrauterine titanium tandem of varying standard lengths x6- compatible with 2.3.2
2.3.5 | CT/MRI titanium vaginal ring applicator set (300, 450and 600), intrauterine tandem (varying lengths) with perineal bar x6
2.3.6 | Rectal retractor x6(compatible with 2.3.5)
2.3.7 | Intrauterine Y type titanium applicator x2
2.3.8 | Vaginal cylinder set with applicator and varying diameter caps including standard accessories x4
2.3.9 | Standard vaginal multichannel applicator with accessories x2
2.3.10 | Surface applicator slabs with minimum 5 channels.
2.3.11 | Universal applicator with x-ray marker plug: channel indicator x4
2.3.12 | Applicator Stabiliser/Jig
2.3.13 | Electronic Mobile Adjustable (x,y,z axis) and electric hand controlled Brachytherapy Radio-Translucent Couch Lithotomy positioning accessories. Minimum weight capacity of 200kg. Must include patient leg support during procedure.
2.3.14 | Oesophageal applicator set varying tube dimensions as per supplier availability x1
2.3.15 | Complete high dose rate prostate brachytherapy solution x1
2.3.16 | Consumable prostate Interstitial tube on contract
2.3.17 | Sterilization Oven or alternate system available from the appointed supplier.
2.4 | Compulsory Dosimetry and Patient Verification equipment
2.4.1 | The following equipment are essential for ensuring quality and safety of treatment and must be included upon delivery:
2.4.2 | Room dose monitoring detector unit with an external electric buzzer, 2 way Intercom system, two in- room zoomable cameras with external monitors (x2).
2.4.3 | Standard absolute dosimetry calibration kit with well ionisation chamber, electrometer, thermometer and barometer with calibration certificate.
2.4.4 | Radiation warning light system (inside and outside the room).
2.4.5 | Check source ruler for daily source position check, with transfer tube.
2.4.6 | Two Boxes of EBT 3 films for Autoradiography.
2.4.7 | Survey meter with external Geiger-Mueller counter.
2.4.8 | Radiation pocket dosimeter
2.4.9 | Independent inroom radiation monitor (ionization chamber)
2.4.10 | Laptop/PC with recent software for Calibration of source activity (Microsoft office, Invivo dosimtery).
2.4.11 | Supply and install a single LED head ceiling mounted theatre light (inclusive on the maintenance warranty) for brachytherapy procedure.
2.5 | Important additional Items to be considered
2.5.1 | Treatment delivery and treatment planning printer
2.5.2 | Data backup solution
2.5.3 | Nasopharynx applicator set.
2.5.4 | Oesophagus applicator set. Size: from smallest to a diameter of 12 mm
2.5.5 | Bronchus applicator set (2 mm diameter)
2.5.6 | Interstitial ring and IU applicator set (CT/MRI)
2.5.7 | Appropriate source insert for well chamber calibration (If CO-60 or X-ray source)
2.6 | POWER SUPPLY
2.6.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
2.6.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
2.6.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
2.6.4 | A 3 phase in / out power supply must be offered. State details.
2.6.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
2.6.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
2.6.7 | The lifetime of UPS must be at least 15 years.
2.6.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
2.6.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
2.6.10 | The mains cable of the unit tendered for must be SABS colour coded. State details.
2.6.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
2.6.12 | Maintenance free batteries must be offered.
2.7 | Brachytherapy C-Arm Table (optional extra)
2.7.1 | Minimum Dimensions: 214.6 x 64.8 cm with OR accessory rails
2.7.2 | Minimum Tabletop: 134.6 x 53.3 cm; 213.4 cm with extension
2.7.3 | Minimum Lightweight Extension (For patient transfer only, not for imaging.): 67.3 x 53.3 cm
2.7.4 | Minimum Radiolucent Extension: 78.7 x 64.8 cm with OR accessory rails
2.7.5 | Minimum Radiolucent Area: 137.2 x 44.5 cm
2.7.6 | Tabletop Material: Three Fiberesin™ sections – each with individual 5cm mattress
2.7.7 | Mattress: Individual, seamless, thick (5 cm), one for each section
2.7.8 | Motions:
2.7.9 | Height Adjustable: 86.4 to 110.5 cm
2.7.10 | Tabletop X Motion (head-to-toe): atleast 50 cm
2.7.11 | Tabletop Y Motion (side-to-side): atleast 25 cm
2.7.12 | Trendelenburg: 0° to 20°
2.7.13 | Reverse Trendelenburg: 0° to 20°
2.7.14 | Lateral Roll: 0° to 20°
2.7.15 | Fowler Positioning: Adjusts from 0° to 80° with touch handle lock/unlock controls
2.7.16 | Hand Control: activates all motions
2.7.17 | Foot Control: activates all motions
2.7.18 | Optional:
2.7.19 | X-Ray Grid: Removable X-Ray grid, 103 line, 8:1 ratio
2.7.20 | Accessory Rails: Standard OR accessory rails 2.86 x .95 cm on table (rails also on table imaging extension)
2.7.21 | Patient Restraints: Two body straps, two ankle straps
2.7.22 | Finish: Powder coat
2.7.23 | Gynaecological position or Lithotomy position: 220 kg patient capacity and use with radiolucent extension.
2.7.24 | Weight tested to four times the patient load rating.
2.7.25 | Power: 115 VAC or 230 VAC and Battery. Battery automatically charges when table is plugged in. An additional battery can be charged with optional wall-mounted charger.
2.7.26 | Certifications: ETL and cETL Listed to UL 60601-1®, CAN/CSA© C22.2