## Sheet1
RADIOTHERAPY TREATMENT PLANNING SYSTEMS
Item | TECHNICAL SPECIFICATIONS | Weight
1 | Treatment Planning system
1.1 | The systems must be compatible with the existing unit in the purchasing institution
1.2 | 2D, 3D, IMRT and VMAT Software License
1.3 | Includes the following features:
1.3.1 | Contouring
1.3.2 | CT Simulation
1.3.3 | Fusion 4D
1.3.4 | Support Plan Review and Approval
1.3.5 | DICOM Import and Export Capabilities
1.3.6 | 3D Planning:
1.3.7 | Support for MLC shapes, blocks, apertures, wedges and electron applicators
1.3.8 | Support for static and dynamic gantry delivery techniques
1.3.9 | Collapsed Cone Algorithm
1.3.10 | Monte Carlo Algorithm for MLC based planning
2 | IMRT software license for static gantry IMRT planning including:
2.1 | Step and Shoot delivery mode, supported for all MLC Types
2.2 | Dynamic MLC (Sweeping Window) delivery mode, supported for Elekta and Varian MLC Types
2.3 | Dynamic conformal Arc with MLC shaping based on the view of the target. Supported for all MLC Types
2.4 | VMAT software license for volume-modulated arc planning including:
2.5 | Fully dynamic gantry and MLC movement
2.6 | Variable dose rate delivery, controlled by varying gantry speed, leaf speed, and dose rate
2.7 | Printer compatible with treatment planning system.
3 | Minimum PC Specification
3.1 | Hardware Package consisting of the following components:
3.2 | Please note that Supplier must provide the latest technology of computer hardware available at time of shipping.
3.2.1 | Base PC - HP Z8 G4 (Dual Processor)
3.2.2 | Processor (x2) - Xeon 14C
3.2.3 | Graphics - NVIDIA Quadro P2000 5.0GB
3.2.4 | Hard drive (x2) Int. SATA SSD 1TB
3.2.5 | Keyboard & Mouse
4 | Contouring systems
4.1 | Software License includes the following features:
4.1.1 | Contouring
4.1.2 | CT Simulation
4.1.3 | Fusion
4.1.4 | 4D Support
4.1.5 | Plan Review and Approval
4.1.6 | DICOM Import and Export Capabilities
4.1.7 | Use with Treatment Planning Systems for dose calculation
5 | Recommended hardware
5.1 | Xeon 4-Core Processor
5.2 | 8GB Memory
5.3 | NVIDIA Video Card
5.4 | 2560 x 1600 Resolution
5.5 | 500GB SATA 6Gb/s Hard Disk
6 | Record and verify system
6.1 | R & V must be Image-Enabled Electronic Charting features required for clinical modules with site licenses.
6.2 | Procedure Management: Advanced procedure management system including complete charge/code capture, audit, and export capability. Site license.
6.3 | IQ Server and SQL. License included. One server is required per department.
6.4 | Image-Guided Treatment Management (IGTM) system that includes support for automated patient set up, image registration and analysis tools to assist the user in evaluating the quality of patient setup based on quantitative comparisons between the reference image and the port film. Site license.
6.5 | Advanced automation and process improvement to deliver greater control, efficiency and customization across your entire oncology workflow. Site license.
6.6 | Import and Export
7 | Oncology Maganement server
7.1 | Server complatitible with data storage for the system used for treatment delivery including record and verify system.
7.2 | Latest model Hardware, software and calculation licences.
7.3 | Server Rack must be included.
8 | POWER SUPPLY
8.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
8.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
8.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
8.4 | A 3 phase in / out power supply must be offered. State details.
8.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
8.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
8.7 | The lifetime of UPS must be at least 15 years.
8.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
8.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
8.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
8.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
8.12 | Maintenance free batteries must be offered.