## Sheet1
DIGITAL FIXED CHEST X-RAY UNIT, CEILING SUSPENDED X-RAY TUBE AND ERECT BUCKY STAND WITH FIXED DETECTOR
ITEM | TECHNICAL SPECIFICATIONS
1.1 | The system must be capable of performing all general radiographic examinations of adults and children
1.2 | The system must be configured with the x-ray tube affixed to a ceiling suspended tube assembly and a vertical buckey stand combination.
1.3 | The system must incorporate one flat panel digital detector that must fit into the erect buckey : 43cm x 43 cm field size range .
1.4 | The system must be capable to capture a "low- energy"image and a "high-energy" image during a single examination. (Dual energy substraction chest imaging)
1.6 | System include LCD 19 inch sized, flat screen colour monitor for operators console /acquisition review station
1.7 | UPS provided to be mounted on a bracket/cabinet, not to be placed directly on the floor.
1.8 | The system must have removable grids.
1.9 | A wall mounted hanger/cabinet for the storage of removable grids and detector must be provided .
2 | X-RAY GENERATOR
2.1 | Generator should be of latest technology with high frequency and high speed starter and a power output of 80 kW and above range. Details are stated in supporting documents provided.
2.2 | A monitoring and display for heat status must be provided.
2.3 | AEC Functionality with 3-field ionization chamber
2.4 | APR Functionality (Anatomical programming) . Details are stated in supporting documents provided
2.5 | Maximum mA Range of 800mA and above at 100kV. Maximum mA output at 100 kV according to IEC60601-2-7 is stated.
2.6 | Selectable kV range of 40-125 kVp in 1 kVp increments
2.7 | The system must have selectable mAs. Details are stated.
2.8 | Shortest exposure time with AEC. Details are stated.
2.9 | Automatic mains compensation capability must be included.
3 | X-RAY TUBE HOUSING AFFIXED TO CEILING SUSPENSION: CEILING SUSPENDED TUBE ASSEMBLY
3.1 | X-ray tube should be a rotating anode high speed, compatible with flat panel detector. X-ray tube must be affixed to the ceiling suspended tube support .Details including latest type of mode, are stated in supporting documents provided.
3.2 | Dual focal spot (small and large focus sizes) must be provided. Details of nominal focal spot sizes and focal spot combinations ,according to IEC 60336 prescription, are stated in supporting documents provided.
3.3 | Real time of tube overloading should be displayed
3.4 | Anode heat storage capacity must be 300 kHU and above. Details are stated in supporting documents provided.
3.5 | Tube overload protection is required
3.6 | Total Filtration at 70kV is 2.5mm (IEC 60601-1-3 Standard)
3.7 | Manual multileaf collimator with rotating flange collimator
3.8 | ± 45° swivelling/rotation of collimator
3.9 | Light localizer on/off switching
3.10. | Collimator must have electronic timer to terminate the light.
3.11 | X-ray Tube rotation range between ± 120 ° up to ± 135 ° . (± symbol means in both directions)
3.12 | Variable SID of the system
3.13 | A ceiling suspended x-ray tube support with a telescopic column must be offered.
3.14 | The ceiling syspended x-ray tube system must have a transverse travel range. (3000mm to 4000mm). Details of range are stated.
3.15 | A ceiling suspended x-ray tube support with a telescopic column must be offered.
3.16 | A centering device in longitudinal and transversal directions to be provided.
3.17 | The vertical travel of the column must not be less than 50cm from the floor. State range of offered system.
3.18 | Electromagnetic brakes with fully counter balanced mechanism
3.19 | Interlocking mechanism for 100 cm/115cm and 180 cm SID . These distances must be electronically displayed on the x-ray tube /collimator housing.
3.20. | A solid control arm with integrated handles as well as an electronic digital angle display panel and individual brake release buttons including a single button releasing all locks to be offered.
3.21. | The digital display panel must also indicate the SID (source image distance) of the unit from the bucky table.
4 | VERTICAL STAND WITH DETECTOR UNIT
4.1 | Vertical stand fitted with a digital flat panel - detector and three AEC measuring chambers
4.2 | Built in tray/support for digital detector and removable grid
4.3 | The digital flat panel detector must be fixed.
4.5 | Vertical height of stand must be 2 metres and above
4.6 | Range of vertical movements of buckey: ( 1) 35cm-45 cm (minimum range height, and (2) 180cm- 200cm (maximum range height).Details are stated
4.7 | Patient hand grips for arms above head and on side of detector holder
5 | DIGITAL FLAT PANEL DETECTOR SYSTEM
5.1 | One x Digital FP detector of 43cm x 43 cm field size.
5.2 | Amorphous silicon with cesium iodide scintillator
5.3 | Detectors have acquisition depth of 14 bits and above grayscale. Details are stated in supporting documents provided.
5.4 | Pixel pitch up to 160 micrometer (µm).
5.5 | Detector to have matrix in pixels of 2000 x 2000 and above
5.7 | DQE must be between a range of 60%-70% @ 0 line pairs per millimeter / 50%- 60 % @ 0.05 line pairs per millimeter / 40 %-50% @ 1 line pairs per millimeter (lp/mm). DQE and MTF details are stated in supporting documents.
5.8 | Description of the scattered radiation grid are detailed in supporting documents provided.
5.9 | The data transmission/ reading speed from exposure to preview is 10 seconds and lower .
5.10. | The data transmission/ reading speed from exposure to full image is 15 seconds and lower .
5.11 | The flat panel battery:
5.11.1 | The battery should be able to acquire and transfer at least 90 images or more when fully charged. Details are stated in the supporting documents provided .
5.11.2 | The battery should be able to handle 5 hours of active use and 15 hours standby, when fully charged. Details are stated in the supporting documents provided .
5.11.3 | Charging time for each battery from fully depleted to fully charged should be less than 2 hours. Details are stated in the supporting documents provided.
5.11.4 | The batteries must be interchangeable between the 24 x30 cm and 43 x 43 cm panels.
5.11.5 | The battery charger must be light weight and should weigh less than 2 kg. Details of the weight of the battery charger are stated in the supporting documents provided .
5.11.6 | The battery must be able to operate at ambient temperature range of minumum of 5°C to maximum of 35 °C.
5.11.7 | The battery must be able to operate at ambient humidity of minimum humidity of 20 % to maximum of 85% RH.
5.11.8 | The battery must be able tolerate storage at a temperature range of minumum of -30 to to maximum of 60 degrees celsius.
5.11.9 | The battery must be able tolerate storage at a minimum humidity of 10 to maximum of 85% RH.
5.12 | The Grid
5.12.1 | An encasement with a built in grid that fits the 43cm x 43cm detector must be supplied.
5.12.2 | The grid must have at least:
5.12.3 | Grid ratio 8:1
5.12.4 | 100cm to 180cm focal distance
5.13 | Load capacity
5.13.1 | The uniform load that the detector must be able to withstand should be at least 150kg. Details are stated in the supporting documents provided .
5.13.2 | The local load (on an area of 40 mm diameter) that the detector must be able to withstand should be at least 100kg. Details are stated in the supporting documents provided .
6 | OPERATOR CONSOLE
6.1 | High resolution digital radiography system with DICOM network connection for image processing and image display on a preview monitor
6.2 | A diagnostic radiology LCD 19 inch imaging workstation based on latest flat panel detector technology for high image dynamics with excellent signal/noise ratio is required
6.3 | Application and examination database sotfware included.Details are stated
6.4 | Image processing software are included
6.5 | 64 bit 3.2GHz i5 Core Dual Microprocessor or similar
6.6 | RAM: 16-32 GB
6.7 | 64 bit Operating system included (Windows or Similar)
6.8 | Image storage capacity up to 10 000 full format images on internal hard drive. Details of image storage capacity are stated in GB, as well as in maximum number of images in the supporting documents provided.
6.9 | No Autodeletion of images stored on internal hard drive. System to alert when near maximum storage capacity is reached.
6.10. | System must be DICOM compatible and capable of HIS/HMS/ RIS and PACS integration
6.11 | DVD /CD Writer inclusive of software for functionality of the DVD/CD Writer..
6.12 | Virus free software
6.13 | Ergonomic keyboard with mouse and function buttons
6.14 | Memory stick support/USB PORT for QC
6.15 | Touch screen monitor
7 | USER INTERFACE
7.1 | Patient data entering
7.2 | Menu
7.3 | Integrated generator operation
7.4 | Menu control
7.5 | Organ program selection
7.6 | Window position/width
7.7 | Horizontal/vertical image mirroring
7.8 | Image manipulation: rotation, cropping, contrast,brightness,sharpness, inversion
7.9 | Left/right, ap/pa working
7.1 | Configurable text annotation
7.11 | Filter selection for edge enhancement
7.12 | Filter selection for noise suppression
7.13 | Image pan and zoom capability
7.14 | measurement tools
8 | CONTROL CONSOLE
8.1 | Patient data entering
8.2 | Manual
8.3 | Bar code reader
8.4 | Monitor size of 19 inch(diagonal) Touch screen
8.5 | Colour
8.6 | Keyboard with mouse and function buttons
8.7 | Spill proof keyboard design
8.8 | Mouse
8.9 | Menu
8.10. | Integrated generator operation
8.11 | Menu control
8.12 | Organ program selection
8.13 | Window position/width
8.14 | Horizontal/vertical image mirroring
8.15 | Image rotation
8.16 | Left/right, ap/pa working
8.17 | Configurable text annotation
8.18 | Filter selection for edge enhancement
8.19 | Filter selection for noise suppression
8.20. | Image zoom capability
9 | APPLICATION AND EVALUATION PROGRAMS
9.1 | Organ-related application program
9.2 | Post processing functionality
9.3 | Windowing (centre/width)
9.4 | Paging forward and back
9.5 | Multi-image display on the monitor (mosaic)
9.6 | Retro-collimation facilities (shutter)
9.7 | Image documentation in the background
9.8 | Multi-tasking technique
9.9 | Edge enhancement
9.1 | Noise suppression
9.11 | Zoom
9.12 | Individual energy substraction programs provided
10 | SYSTEM INTERPHASE
10.1 | Detector interphase
10.2 | Interface DICOM: "Basic Print"
10.3 | Interface DICOM: "Modality Worklist"
10.4 | Interface DICOM: "Storage"
10.5 | Interface DICOM: "Send as SCU"
10.6 | Interface DICOM "Image export/transfer to PACS/HIS/RIS
10.7 | Is there a "Confirmation Statement" for all DICOM interfaces
10.8 | Firewall protection
10.9 | The succesful bidder must ensure that the configuration between the processed images on the Digital Imaging Workstation and that transferred to the receivers of these images, such as PACS , Laser Printer , CD/DVD are properly calibrated/ reconfigured during installation, service and maintenance,repairs, and future software updates/upgrades .
11 | OPTIONAL ITEMS
11.1 | 1 X TB HARD DRIVE(external) ; DICOM CONFIGURED TO STORE, TRANSFER AND PRINT X-RAY IMAGES.
11.2 | Motorized collimation control as an option
11.3 | Tilting buckey as an option
11.4 | Synchronized tube and detector movements. Auto detector tracking and autotube positioning
11.5 | The system must have the ability to perform digital long length stitching images(scoliosis/long bones) -OPTIONAL
11.6 | Auto image stitching program included in the software package.
11.7 | TOR Rad or similar Image Resolution QC Test Tool
11.8 | A trolley with lockable castors, radiolucent top and an integrated detector holder that must move across the length of the trolley.
12 | RADIATION PROTECTION
12.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
12.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
12.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
12.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
12.5 | A radiation report must be automatically generated and transferable via DICOM.
12.6 | Should be FDA or CE approved product
12.7 | Electrical safety conforms to standards for electrical safety IEC-60601
12.8 | The unit must comply to IEC-61267
12.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
12.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
12.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
12.12 | Lead Glass window to be installed as per room requirement
12.13 | Safety aspects of Radiation dosage leakage should be spelt out
12.14 | DB boards and electrical filters to be installed
12.15 | QA TOOLS
12.16 | Display Exposure index values
12.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
12.18 | Patient Conventional face shield to be included.
12.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
12.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
12.21 | Interlocks to stop system movement when compression is applied must be provided.
12.22 | Interlock for premature release of the X-ray switch must be available
12.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
12.24 | Phantom for calibration should be provided.
13 | POWER SUPPLY
13.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
13.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
13.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
13.4 | A 3 phase in / out power supply must be offered. State details.
13.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
13.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
13.7 | The lifetime of UPS must be at least 15 years.
13.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
13.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
13.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
13.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
13.12 | Maintenance free batteries must be offered.
14 | AIR CONDITIONER
14.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
14.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
14.3 | The air conditioner must also be included in the maintenance plan.
14.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
14.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
14.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
14.7 | Air-conditioning must be included in Maintenance plan