## Sheet1
DIGITAL MOBILE X-RAY UNIT
Item | SPECIFICATIONS
1.1 | The unit must be a light weight, motorized, compact system and easy to manoeuvre.
1.2 | The unit must have a built- in battery.
1.3 | The unit must have a lock function.
1.4 | The unit must have electro-magnetic locks.
1.5 | Tube rotation must be 180 degrees around the vertical axis.
1.6 | A manual collimator with +90 degrees and –90 degrees swivel must be included.
1.7 | “All free” lock control buttons on the collimator handle must be available for fast positioning, including the Film Focus Distance (FFD) setting, tube rotation, extension and contraction of the cross-arm.
1.8 | The unit's focus to floor distance should at least be 1500 mm.
1.9 | The unit must easily pass through a standard doorway (at least 800 mm width, height 2100mm).
1.10 | The unit must have a lead apron hanger.
1.11 | The unit offered must be able to make sharp spot turning.
1.12 | Large sturdy wheels should be incorporated.
1.13 | The unit should have a tube storage compartment in front of the control panel for safe transport.
2 | DIGITAL MOBILE FLAT PANEL DETECTOR WITH IMAGE PROCESSING
2.1 | The detectors must be of the flat panel type and wireless.
2.2 | 2 flat panel detectors must be provided (sizes at least 35x43cm for adults & 24x30cm for children).
2.3 | Detectors must be of Cesium Iodide (Csl) material.
2.4 | The pixel size (matrix size) must be not > 150µm
2.5 | The acquisition depth must at least be 12 bit.
2.6 | Weight of the detector panel must not be more than 5kg including the battery.
2.7 | The detectors must be supplied with 1 additional battery.
2.8 | The unit must have detector storage compartments for various sizes
2.9 | The charging station for the batteries must be supplied.
2.10 | A grid for the wireless detectors must be provided
2.11 | The flat panel battery:
2.11.1 | The battery should be able to acquire and transfer at least 90 images or more when fully charged. Details are stated in the supporting documents provided .
2.11.2 | The battery should be able to handle 5 hours of active use and 15 hours standby, when fully charged. Details are stated in the supporting documents provided .
2.11.3 | Charging time for each battery from fully depleted to fully charged should be less than 2 hours. Details are stated in the supporting documents provided.
2.11.4 | The batteries must be interchangeable between the 24 x30 cm and 43 x 43 cm panels.
2.11.5 | The battery charger must be light weight and should weigh less than 2 kg. Details of the weight of the battery charger are stated in the supporting documents provided .
2.11.6 | The battery must be able to operate at ambient temperature range of minumum of 5°C to maximum of 35 °C.
2.11.7 | The battery must be able to operate at ambient humidity of minimum humidity of 20 % to maximum of 85% RH.
2.11.8 | The battery must be able tolerate storage at a temperature range of minumum of -30 to to maximum of 60 degrees celsius.
2.11.9 | The battery must be able tolerate storage at a minimum humidity of 10 to maximum of 85% RH.
2.12 | The Grid
2.12.1 | An encasement with a built in grid that fits the 43cm x 43cm detector must be supplied.
2.12.2 | The grid must have at least:
2.12.3 | Grid ratio 8:1
2.12.4 | 100cm to 180cm focal distance
2.13 | Load capacity
2.13.1 | The uniform load that the detector must be able to withstand should be at least 150kg. Details are stated in the supporting documents provided .
2.13.2 | The local load (on an area of 40 mm diameter) that the detector must be able to withstand should be at least 100kg. Details are stated in the supporting documents provided .
2.14 | Reference image to be displayed on the mobile monitor within at least 3 seconds.
2.15 | Storage of not less than 2,000 images must be possible.
2.16 | DICOM format transfer to an image server and/or laser printer must be possible.
2.17 | After exposure, various types of processing must be possible to obtain the required image information selectively.
3 | X-Ray Generator (Micro-Processor Controlled High Frequency)
3.1 | The nominal kilowatt output rating may not be less than 30 kW.
3.2 | The system must be single phase( 230 V, 50Hz) system.
3.3 | The unit should have anatomical programming facilities with at least 300 programs.
3.4 | Automatic overload protection must be standard on the generator with error code readout.
3.5 | Automatic mains compensation must be standard on the system (State the tolerance of mains power fluctuation permissible on the system).
3.6 | Minimum exposure time may not be greater than 1 millisecond.
3.7 | The kV range must be from 40 kV (minimum) to 125 kV (maximum).
3.8 | At 60 kV the unit must be able to give at least 320 mA.
3.9 | The generator must have a high speed starter unit to drive the x-ray tube to at least 3000 rpm.
4 | X-Ray Tube and Cables
4.1 | The unit must have high speed anode rotation with dual focus x-ray tube.
4.2 | The anode rotation speed must be at least 3000 rpm.
4.3 | The unit must have focus sizes of at least 0,6 mm (small) and 1,2 mm (large).
4.4 | A full size multileaf collimator unit with integrated centering lamp must be offered. Double-slot collimator units are not acceptable.
4.5 | A rotation flange must be provided with the unit for full rotation of the collimator on the x-ray tube.
4.6 | Minimum anode heat storage capacity must be at least 100 kHU.
5 | Control Console
5.1 | The control console must have the following patient data entering functions:
5.1.1 | Automatically retrieve patient data via LAN and Wireless network
5.1.2 | Manual patient registration must be possible
5.2 | Touch panel LCD monitor or latest technology.
5.3 | Monitor size must be at least 15 inch
5.4 | Image processing must be possible
5.5 | The Menu Control should be able to do the following:
5.5.1 | Organ program selection
5.5.2 | Window position/width
5.5.3 | Horizontal and vertical image mirroring
5.5.4 | Image rotation
5.5.5 | Antero-Posterior/ Postero-Anterior views
5.5.6 | Configurable text annotation
5.5.7 | Filter selection
5.5.8 | Image zoom
5.5.9 | Must be able to print different layouts of selected images.
5.5.10 | Edge enhancement and noise suppression
5.5.11 | Quality management program including reject analysis
5 | RADIATION PROTECTION
5.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
5.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
5.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
5.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
5.5 | A radiation report must be automatically generated and transferable via DICOM.
5.6 | Should be FDA or CE approved product
5.7 | Electrical safety conforms to standards for electrical safety IEC-60601
5.8 | The unit must comply to IEC-61267
5.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
5.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
5.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
5.12 | Lead Glass window to be installed as per room requirement
5.13 | Safety aspects of Radiation dosage leakage should be spelt out
5.14 | DB boards and electrical filters to be installed
5.15 | QA TOOLS
5.16 | Display Exposure index values
5.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
5.18 | Patient Conventional face shield to be included.
5.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
5.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
5.21 | Interlocks to stop system movement when compression is applied must be provided.
5.22 | Interlock for premature release of the X-ray switch must be available
5.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
5.24 | Phantom for calibration should be provided.
6 | POWER SUPPLY
6.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
6.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
6.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
6.4 | A 3 phase in / out power supply must be offered. State details.
6.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
6.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
6.7 | The lifetime of UPS must be at least 15 years.
6.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
6.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
6.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
6.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
6.12 | Maintenance free batteries must be offered.
7 | AIR CONDITIONER
7.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
7.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
7.3 | The air conditioner must also be included in the maintenance plan.
7.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
7.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
7.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
7.7 | Air-conditioning must be included in Maintenance plan