## Sheet1
DIGITAL PANORAMIC DENTAL X-RAY UNIT
Item | SPECIFICATIONS | Weight
1.1 | The Unit must be a Digital Panoramic Dental X - ray Unit | #
1.2 | The unit must have a software system capable of automatically removing redundant tomographic shadows and ghost images that result in geometric unsharpness, e.g. ghost images/shadows of the cervical spine and the occiput. | #
1.3 | The unit must allow for optimum selection of exposure factors to reduce the radiation dose. | #
1.4 | It must be capable of performing image segmenting to reduce radiation, e.g. Vertical segmenting to expose a smaller area of diagnostic interest. | #
1.5 | The unit must have an automatic doubleTemporo - Mandibular Joint ( TMJ) program to perform full temporo-mandibular studies on a single display or film. | #
1.6 | The unit must have a sinus program, which will show a clear view of the maxillary sinuses. | #
1.7 | The GUI must be simple and intuitive. Text and or graphic symbols should prompt the operator. | #
1.8 | All necessary information must be displayed on the main console. | #
1.9 | The unit must be equipped with a self-diagnostic test system, capable of displaying error messages/codes shall be incorporated into the unit. | #
1.1 | The unit must have a fixed bite block which is easy to clean after each patient. | #
1.11 | The unit must incorporate a fully automatic film marking system for panoramic radiography to mark the exposure factors, the chosen program, as well as patient identification information | #
1.12 | The unit must have a stable headrest for temple support | #
1.13 | The rating of the x ray tube must be at least 2.5kW.Please supply rating charts. | #
1.14 | The tube focal spot size must range from 0,5 X 0,5 mm to 0.8x0.8mm. | #
1.15 | The total tube filtration must be 2.5mm AI. | #
1.16 | The anode voltage must be from 60kVp to 90 kVp. | #
1.17 | The tube anode current to be from 4 to 10 mA or better | #
1.18 | The unit must have exposure time of 4.5 to 20 seconds continuous or split. | #
1.19 | The unit must include an electronic kV compensation over the cervical spine. | #
1.2 | Exposure factors must be pre-programmed, and also manually selectable. | #
2 | SYSTEM FUNCTIONALITY
2.1 | The Unit must have fully customizable patient work list
2.2 | The Unit must have examination related pre-configured views
2.3 | System must set automatically all generator settings
2.4 | System must set collimators automatically
2.5 | System must set pre-filters automatically
2.6 | Manual input for patient data must be possible (Please describe)
3 | WORKSTATION
3.1 | A WINDOWS Operating computer system is required with at least 20 GB hard drive, a network card and suitable graphics card. (State capacity)
3.2 | At least a 17 inch touch screen flat panel LCD monitor (supporting at least 3 Megapixel resolution) must be supplied.
3.4 | A Mono Laser jet printer with at least 1200 dpi must be provided
4 | IT Integration
4.1 | The system must have the following DICOM compatibility:
4.1.1 | DICOM print class
4.1.2 | DICOM storage class
4.1.3 | DICOM query/retrieve
4.1.4 | DICOM worklist
4.1.5 | The system must comply to the following IHE profiles (Integrating the Healthcare Enterprise):
4.1.6 | Scheduled Workflow (SWF)
4.1.7 | Simple image and Numeric Reports (SINR)
4.1.8 | Patient Information Reconciliation (PIR)
5 | RADIATION PROTECTION
5.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
5.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
5.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
5.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
5.5 | A radiation report must be automatically generated and transferable via DICOM.
5.6 | Should be FDA or CE approved product
5.7 | Electrical safety conforms to standards for electrical safety IEC-60601
5.8 | The unit must comply to IEC-61267
5.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
5.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
5.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
5.12 | Lead Glass window to be installed as per room requirement
5.13 | Safety aspects of Radiation dosage leakage should be spelt out
5.14 | DB boards and electrical filters to be installed
5.15 | QA TOOLS
5.16 | Display Exposure index values
5.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
5.18 | Patient Conventional face shield to be included.
5.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
5.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
5.21 | Interlocks to stop system movement when compression is applied must be provided.
5.22 | Interlock for premature release of the X-ray switch must be available
5.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
5.24 | Phantom for calibration should be provided.
6 | POWER SUPPLY
6.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
6.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
6.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
6.4 | A 3 phase in / out power supply must be offered. State details.
6.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
6.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
6.7 | The lifetime of UPS must be at least 15 years.
6.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
6.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
6.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
6.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
6.12 | Maintenance free batteries must be offered.
7 | AIR CONDITIONER
7.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
7.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
7.3 | The air conditioner must also be included in the maintenance plan.
7.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
7.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
7.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
7.7 | Air-conditioning must be included in Maintenance plan