## Sheet1
DIGITAL X-RAY UNIT, U-ARM TYPE
ITEM | TECHNICAL SPECIFICATIONS
1 | CEILING SUSPENDED / MOUNTED X-RAY UNIT, U-ARM TYPE, INDEPENDENT FOR A SINGLE BAY ( IN TRAUMA UNIT)
1.1 | Counter balance system.
1.2 | One cordless / wireless detector .
1.3 | One workstation with a viewing station.
1.4 | Ease of movement to be well secured with locks at the tube head, if possible, or at the arm. Details must be stated.
1.5 | Cables to be well secured, avoid dangling above operators heads.
1.6 | Horizontal / transverse travel to be 3 meters and more, as per need in area positioned.
1.7 | The vertical travel of the U-arm must reach as low as 40-60cm above the floor.
1.8 | Distance from the detector to the X-Ray source (SID) must not be less than 100cm.
1.9 | Distance from the detector to the X-Ray source (SID) must be adjustable.
1.1 | The projection direction for the recumbent patients must include all angles between +90° and -90°.
1.11 | The X-Ray system must be equipped with collision protection.
2 | ONE DETECTOR TRAY
2.1 | The tube on the one end of the U-arm and the detector must be synchronized.
2.2 | The detector tray must be integrated with the U-arm.
2.3 | A battery pack must be provided to charge the detector.
2.4 | Tether interphase cable, LAN cable, generator interphase cables must be provided.
2.5 | The detector tray must accommodate both the detector and the removable grid.
2.6 | The grid ratio must be 10:1 lines per centimeter. The details must be stated.
2.7 | 43cm x 43cm Detector tray.
3 | ONE FLAT PANEL DETECTOR
3.1 | Pixel pitch up to 160 micrometer (µm).
3.2 | The flat panel detector must be wireless/fixed
3.3 | Detectors have acquisition depth of 14bits and above grayscale. Details are stated in supporting documents provided.
3.4 | The size of the flat panel detector must be between 35cm x 43cm to 43cm x 43cm.
3.5 | Detector to have matrix in pixels of 2000 x 2000 and above
3.6 | Patient’s information/ image preview data must be transferable in not more than 3 seconds.
3.7 | The flat panel detector must be DICOM compliant and enabled.
3.8 | The weight of the the flat panel detector must be less than 5kg.
3.9 | The flat panel detector must be of Cesium Iodide material.
3.10. | DQE must be between a range of 60%-70% @ 0 line pairs per millimeter / 50%- 60 % @ 0.05 line pairs per millimeter / 40 %-50% @ 1 line pairs per millimeter (lp/mm). DQE and MTF details are stated in supporting documents.
3.11 | The wall mounted cabinet must be installed for storage of removable wireless detector (optional) , QC test Tools and Grid.
3.12 | The flat panel battery:
3.12.1 | The battery should be able to acquire and transfer at least 90 images or more when fully charged. Details are stated in the supporting documents provided .
3.12.2 | The battery should be able to handle 5 hours of active use and 15 hours standby, when fully charged. Details are stated in the supporting documents provided .
3.12.3 | Charging time for each battery from fully depleted to fully charged should be less than 2 hours. Details are stated in the supporting documents provided.
3.12.4 | The batteries must be interchangeable between the 24 x30 cm and 43 x 43 cm panels.
3.12.5 | The battery charger must be light weight and should weigh less than 2 kg. Details of the weight of the battery charger are stated in the supporting documents provided .
3.12.6 | The battery must be able to operate at ambient temperature range of minumum of 5°C to maximum of 35 °C.
3.12.7 | The battery must be able to operate at ambient humidity of minimum humidity of 20 % to maximum of 85% RH.
3.12.8 | The battery must be able tolerate storage at a temperature range of minumum of -30 to to maximum of 60 degrees celsius.
3.12.9 | The battery must be able tolerate storage at a minimum humidity of 10 to maximum of 85% RH.
3.13 | The Grid
3.13.1 | An encasement with a built in grid that fits the 43cm x 43cm detector must be supplied.
3.13.2 | The grid must have at least:
3.13.3 | Grid ratio 8:1
3.13.4 | 100cm to 180cm focal distance
3.14 | Load capacity
3.14.1 | The uniform load that the detector must be able to withstand should be at least 150kg. Details are stated in the supporting documents provided .
3.14.2 | The local load (on an area of 40 mm diameter) that the detector must be able to withstand should be at least 100kg. Details are stated in the supporting documents provided .
4 | ONE WORKSTATION
4.1 | The workstation must be connected to the PACS system, if available at the hospital.
4.2 | The successful bidder must provide the license to connect the workstation to the PACS and link to hospital network.
4.3 | The size of the touch screen LCD monitor must be 19” or above.
4.4 | Enhanced and extensive connectivity with MHL, display port, HDMI, VGA, USB ports.
4.5 | Full HD resolution and wide viewing angle 1000 x 1000 matrix.
4.6 | Stand designed uniquely for the workstation.
4.7 | Keyboard and mouse to be included.
4.8 | Repeat and reject analysis software package must be included.
4.9 | QC software package for viewing monitors must be included (SMPTE pattern or similar.)
4.1 | Browser must be compatible to Windows 10 or higher.
5 | IMAGE ACQUISITION AND PROCESSING
5.1 | Must have an organ related application program.
5.2 | Must be able to configure the organ related application program to individual parameters.
5.3 | Pre- and post-processing of image parameters must be included.
5.4 | The following must be available on image processing:
5.5 | a. Identification of patients
5.6 | b. ID marker selection Left or Right
5.7 | c. AP/PA selection
5.8 | Image functions to include:
a. PAN and zoom capability
b. Rotation
c. Inversion
d. Windowing
e. Filtration
f. Annotation
g. Subtraction
h. Measurement
i. Image manipulation
j. Filter selection for noise suppression and edge enhancement.
5.9 | The image contrast must be adjustable.
5.10. | All updated parameters to be included at no extra cost.
6 | Image documentation and archiving:
6.1 | a. DICOM Print, Transfer, Modality, Work List, Send, Storage and Storage Commit and MPPS.
6.2 | b. All DICOM conformance statements must be provided.
6.3 | c. Storage capacity for 10 000 or more images must be available on the internal hard drive.
7 | ONE X-RAY GENERATOR
7.1 | Must be a three phase, high frequency system.
7.2 | Output must be 80kW and above.
7.3 | The minimum KV must be 40KV up to 125KV at reasonable incremental values, 1 KV step. State incremental step.
7.4 | The mA must be 800mA for a 80kW generator at 100kV. The details must be stated.
7.5 | The generator must have selectable mAs.
7.6 | The settings on the operator console must be displayed digitally.
7.7 | Troubleshooting, overload protection and error codes to be incorporated into the system.
8 | ONE CEILING SUSPENDED X-RAY TUBE ON THE U-ARM
8.1 | Rotating anode.
8.2 | Dual Focal Spot (small and large focus).
8.3 | Anode Heat Storage Capacity must be 300kHU or above.
8.4 | Multi lead collimator to have rotating flange for collimation around the central beam.
8.5 | Swivel / Rotation of collimator around its own axis +/-45 degrees.
8.6 | Light localizer with electronic switch.
9 | CPU HARD-DRIVE
9.1 | Must have a high storage capacity of 1TB.
9.2 | DVD RW and readable facilities to be available.
9.3 | Random access memory must not be less than 4GB.
9.4 | Must have a i5 core Intel processor or higher.
9.5 | Integrated graphics.
9.6 | Antivirus software protector to be provided.
10 | RADIATION PROTECTION
10.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
10.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
10.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
10.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
10.5 | A radiation report must be automatically generated and transferable via DICOM.
10.6 | Should be FDA or CE approved product
10.7 | Electrical safety conforms to standards for electrical safety IEC-60601
10.8 | The unit must comply to IEC-61267
10.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
10.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
10.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
10.12 | Lead Glass window to be installed as per room requirement
10.13 | Safety aspects of Radiation dosage leakage should be spelt out
10.14 | DB boards and electrical filters to be installed
10.15 | QA TOOLS
10.16 | Display Exposure index values
10.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
10.18 | Patient Conventional face shield to be included.
10.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
10.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
10.21 | Interlocks to stop system movement when compression is applied must be provided.
10.22 | Interlock for premature release of the X-ray switch must be available
10.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
10.24 | Phantom for calibration should be provided.
11 | POWER SUPPLY
11.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
11.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
11.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
11.4 | A 3 phase in / out power supply must be offered. State details.
11.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
11.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
11.7 | The lifetime of UPS must be at least 15 years.
11.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
11.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
11.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
11.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
11.12 | Maintenance free batteries must be offered.
12 | AIR CONDITIONER
12.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
12.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
12.3 | The air conditioner must also be included in the maintenance plan.
12.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
12.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
12.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
12.7 | Air-conditioning must be included in Maintenance plan
12 | OPTIONAL ITEMS
12.4 | One LCD monitor of 18" or above.