## Sheet1
Digital X-RAY UNIT (80kW generator)
Description
ITEM
1 | The said Unit must be fully automatic, digital, radiographic examination and evaluation workstation based on two fixed/wireless detectors technology for high image dynamics with excellent signal/noise ratio is required.
1.1 | The bid price quoted must include the supply, installation, demonstration and commissioning of the equipment as well as the building alterations that may be required for the successful installation of the unit/s.
1.2 | State make and model of main unit offered.
1.3 | The bidders are requested to structure their quote so that it reflects the price of the unit separately from any building alterations that are required. The price schedule at the end of this specification must be completed.
1.4 | The unit must be suitable for general and trauma radiography.
2 | GENERATOR:
2.1 | The unit should be a high frequency type generator.
2.2 | The output shall be 80kW. State maximum milliamps at different kilovolts.
2.3 | The unit must incorporate an electronic timer with minimum exposure time of 1mS and electronic switching
2.4 | The unit must include automatic overload protection
2.5 | The unit must have automatic mains compensation.
2.6 | A milliamp meter or indicator must be provided.
2.7 | The kV, mA and time must be independently selectable and digitally displayed.
2.8 | Bidders to state the range of the exposure timer. Acceptable range 0.001 sec – 6 sec (1 ms – 6sec).
2.9 | Bidders to state kV and mA ranges for radiography. The kV range to be at least 40 to 150 kV and mA range at least 10 to 630mA. The kV selection must be in steps of 1 and must be a single control.
2.10. | The unit offered must be programmable. Bidders to supply details of programming facility and the number of programs.
2.11 | The offered unit must include automatic exposure control.
2.12 | Tube overload protection is required
3 | X-Ray TUBE:
3.1 | The rotating anode tube must match the output of the generator.
3.2.1 | The maximum tube kV must be at least 150 kV.
3.2.2 | The unit must incorporate an electronic timer with minimum exposure time of 1mS and electronic switching
3.2.3 | The maximum tube mA must be at least 500 mA
3.2.4 | The speed of anode rotation must be at least 8000rpm
3.2.5 | Anode heat storage capacity must be at least 300kHU
3.2.6 | The tube housing heat storage capacity must be more than 400kHU
3.2.7 | The size of small focus must be 0,6 mm
3.2.8 | The maximum tube mA must be at least 500 mA
3.2.9 | The speed of anode rotation must be at least 8000rpm
3.2.10 | The size of large focus must be 1,2 mm
3.2.11 | Inherent filtration of at least 0.8mmAl
3.2.12 | Total filtration must exceed 2.5mmAl.
3.2.13 | Anode angle must be large enough to provide a uniform beam over a field of 43cm x 43cm at a FFD of 100cm.
3.3 | HIGH TENSION CABLE: One pair of high tension cable of suitable length must be supplied. The cable length must not restrict free movement of the tube for all radiographic procedures.
3.4 | The power supply must be a three phase voltage 380-400 V at 50Hz
4 | X-RAY TUBE WITH MULTILEAF COLLIMATOR
4.1 | Nominal voltage range (40 - 150kV)
4.2 | State the nominal power
4.3 | Focal spot parameter according to International Electrotechnical Commission (IEC) 336
4.4 | Small focus 0.6mm
4.5 | Large focus 1.2mm
4.6 | Anode angle 12o
4.8 | State the speed of rotating anode
4.9 | Heat storage capacity of the anode should not be less than 350kHU
4.10 | Heat storage capacity of the X-ray tube should not be less than 400kHU
4.11 | Radiation absorption (according to IEC 601-1-3)
4.12 | There must be a measuring tape on the X-Ray tube
4.13 | Light localizer on/off switching with timer to terminate the light after 30 seconds
4.14 | The table must be able to perform tomography. Please state the angles offered and the speeds related to each angle.
5 | CEILING MOUNTED TUBE SUPPORT
5.1 | The tube column must be ceiling mounted
5.2 | The X-ray tube assembly rotation must be 360° about the vertical and horizontal axis
5.3 | Tube movements:
5.3.1 | Longitudinal: at least 150 cm from the iso centre
5.3.2 | Transverse: at least 100 cm from the iso centre
5.3.3 | Vertical: Tube centre should be able to move down to the lowest of 30cm above the floor.
5.4 | Multiplanner movements controlled by electric lock
5.5 | All movements to be counter balanced
6 | CEILING SUSPENSION:
6.1 | The ceiling suspension must provide coverage of an area of approximately 3 X 4 meters.
6.2 | Stops must be provided for centering to table and vertical Bucky.
6.3 | Bidders to indicate focal film distance range for vertical and table Bucky.
6.4 | A safety device for possible failure of the counter-balance system must be incorporated.
6.5 | The unit must be able to perform horizontal beam radiography on both side of the table.
6.6 | Bidders to state:
6.6.1 | Tube rotation and movement around vertical axis
6.6.2 | Tube rotation around horizontal axis and
6.6.3 | Range of vertical movement.
7 | BUCKY TABLE:
7.1 | The bidder to state the make and model of the Bucky table offered.
7.2 | The Bucky table must have a flat, four way floating top.
7.3 | The table must be electronically height controlled. Foot switches must form part of the table.
7.4 | State the variable height of the table offered.
7.5 | The dimensions of the table top must have atleast length 210cm and 80 to 90 cm width
7.6 | State the lateral and longitudinal movement of the table top.
7.7 | The table offered must be able to carry patients with a weight of at least 220kg centred and off-centred.
7.8 | State the longitudinal movement of Bucky tray. Bidders must detail whether the distance relates to the outermost end of the Bucky tray, or from centre point to centre point. A data sheet must be included to clarify this point.
7.9 | The table-top movement must be controlled with electromagnetic brakes.
7.10. | The table-top film distance must be approximately 6 centimeters.
7.11 | The Bucky assembly must be fitted with the oscillating secondary radiation grids suitable for direct digital flat panel technology.
7.12 | A compression band with fitting mechanism compatible to the table top must be supplied.
7.13 | A full set of positioning sponges with washable water resistance covers must be supplied with the table.
7.14 | Emergency stop mechanism must be attached to the table
8 | VERTICAL BUCKY
8.1 | State the make and model of vertical Bucky offered.
8.2 | The vertical height should be adequate for a lateral cervical spine projection of a standing patient and therefore maximum height from centre of the Bucky to floor should be at least 180 cm.
8.3 | The Bidder to state range of vertical movement from floor to centre point of the Bucky.
8.4 | Vertical movement of the vertical Bucky must be via an electronic locking/unlocking switch. Analogue locking/unlocking levers or knobs will not be accepted
8.5 | The Bucky assembly must be fitted with the oscillating secondary radiation grids suitable for direct digital flat panel technology and must accommodate a chest radiography at 180cm.
8.6 | Radiography of a patient who is in a seated position on the stretcher or table top must be possible.
8.7 | The Bucky must be mounted centrally on the column with the choice of loading on either left or right.
8.8 | An overhang arm support must be included that is compatible with erect Bucky.
8.9 | An overhang frame for out of Bucky radiography must be included that is compatible with erect Bucky.
9 | DIGITAL FLAT PANEL DETECTOR:
9.1 | The detector technology must be of Cesium Iodide type
9.2 | The exposure formats must be 43cm x 43cm
9.3 | The full vertical rotation must be possible without removing detector for 43 x 43 format
9.4 | The full horizontal rotation must be possible without removing detector for 43 x 43 format
9.5 | The detector exposure matrix: pixels must be minimum 3.6 lp/mm
9.6 | The detector pixel pitch must be minimum 139um
9.7 | The reading speed at maximum format from the exposure to preview on the monitor must not be greater than 8 seconds
9.8 | The digital Flat panels must operate wirelessly
9.9 | The wireless flat panels must be supplied with x2 spare battery packs each that can be charged outside of the flat digital flat panel.
9.1 | The analog to digital conversion must not be less than 12 bits
9.11 | The wireless ditector must come with grids, detector cover and holder
9.12 | The flat panel battery:
9.12.1 | The battery should be able to acquire and transfer at least 90 images or more when fully charged. Details are stated in the supporting documents provided .
9.12.2 | The battery should be able to handle 5 hours of active use and 15 hours standby, when fully charged. Details are stated in the supporting documents provided .
9.12.3 | Charging time for each battery from fully depleted to fully charged should be less than 2 hours. Details are stated in the supporting documents provided.
9.12.4 | The batteries must be interchangeable between the 24 x30 cm and 43 x 43 cm panels.
9.12.5 | The battery charger must be light weight and should weigh less than 2 kg. Details of the weight of the battery charger are stated in the supporting documents provided .
9.12.6 | The battery must be able to operate at ambient temperature range of minumum of 5°C to maximum of 35 °C.
9.12.7 | The battery must be able to operate at ambient humidity of minimum humidity of 20 % to maximum of 85% RH.
9.12.8 | The battery must be able tolerate storage at a temperature range of minumum of -30 to to maximum of 60 degrees celsius.
9.12.9 | The battery must be able tolerate storage at a minimum humidity of 10 to maximum of 85% RH.
9.13 | The Grid
9.13.1 | An encasement with a built in grid that fits the 43cm x 43cm detector must be supplied.
9.13.2 | The grid must have at least:
9.13.3 | Grid ratio 8:1
9.13.4 | 100cm to 180cm focal distance
9.14 | Load capacity
9.14.1 | The uniform load that the detector must be able to withstand should be at least 150kg. Details are stated in the supporting documents provided .
9.14.2 | The local load (on an area of 40 mm diameter) that the detector must be able to withstand should be at least 100kg. Details are stated in the supporting documents provided .
11 | X-RAY EXAMINATION TABLE
11.1 | Table movements must be motorised (all directions)
11.2 | Minimum height (above the Floor) 40-60cm
11.3 | Must have a Floating table top
11.4 | Detector carriage should have a motorized scattered radiation grid.
11.5 | Film Focus Distance = 115cm
11.6 | Table top dimensions of at least 220L X 80W cm.
11.7 | Table top longitudinal movement of at least (50cm) in both direction with minimum patient weight of 220kg.
11.8 | Transversal movement of at least (20cm) in one direction with minimum patient weight of 220kg.
11.9 | Control of the table functions via footswitch attached to the table
11.10 | Electromagnetic brakes to be standard
11.11 | Table top to be of suitable low radiation absorption material
11.12 | Load capacity must be at least 220 kg
11.13 | Must have automatic exposure device
12 | IMAGE WORKSTATION
12.1 | Storage capacity must be minimum 4 tera bytes
12.2 | The system must have minimum of 4 GB memory
12.3 | Minimum 19 inch, maximum 23 inch colour LCD/LED display
12.4 | Must be a touch screen
12.5 | Resolution Minimum 1920x 1080 pixels
13 | IMAGING SOFTWARE
MUST INCLUDE:
13.1 | General Features
13.1.1 | Windows based graphic user interface
13.1.2 | Multi-image display (1x1 / 2x2 / 3x3 / 4x4 )
13.1.3 | Multi-image selection
13.1.4 | Unlimited procedure set ups
13.1.5 | ROI changing and creation feature
13.1.6 | English language support
13.1.7 | Unlimited PACS code ( CPT code )
13.1.8 | Default anatomic program more than 700
13.1.9 | Support DICOM Worklist SCU, DICOM Storage SCU and transfer function
13.1.10 | Support DICOM Multi-transfer function
13.1.11 | All hardware & software relating to DICOM compatibility, including licences, must be supplied.
13.2 | Post processing features
13.2.1 | Noise reduction parameter
13.2.2 | Global contrast parameter
13.2.3 | Detail contrast parameter
13.3 | Image Maintenance (All functions must be supported by the pop-up menu)
13.3.1 | ROI
13.3.2 | Mark
13.3.3 | Horizontal Flip
13.3.4 | Vertical Flip
13.3.5 | Rotate CW
13.3.6 | Rotate CCW
13.3.7 | Inverse (Black or White)
13.3.8 | Text Annotation
13.3.9 | Ruler : Distance tool
13.3.10 | Angle : Angle measurement tool
13.3.11 | Zoom : Image zoom in/out
13.3.12 | Magnify : Image magnify glass window
13.3.13 | Pan : Image panning
13.3.14 | Fit Image : Auto fitting to window size
13.3.15 | Image Cut : Image crop/cut function
13.3.16 | Image Copy : Copy of image in the region of interest(ROI)
13.3.17 | Image Recovery : Recover the original image
13.4 | CD Burning
13.4.1 | DICOMDIR based CDR data generation
13.4.2 | Support CD/DVD Recording
13.4.3 | Include internal DICOM Viewer
13.4.4 | Support multi-study data
13.5 | DICOM Features : DICOM PRINT
13.5.1 | DICOM 3.0 compatible
13.5.2 | Support Print Preview
13.5.3 | Support Film Layout : 1:1 / 1:2 / 2:1 / 2:2
13.5.4 | Support multi-page printing
13.5.5 | Support Real size printing
13.5.6 | Support image swap in layout
13.6 | DICOM Feature : DICOM STORAGE
13.6.1 | DICOM 3.0 compatible
13.6.2 | Support DX/CR modality
13.6.3 | Support the modification of Transfer Syntax
13.6.3.1 | Transfer status
13.6.4 | DICOM 3.0 compatible
13.6.5 | Support DX/CR modality
13.7 | DICOM Feature : WORKLIST
13.7.1 | Support DICOM Modality Worklist Standard
13.7.2 | Support DICOM Query/Retrieve
13.7.3 | Support Search Filter ( ID / Name / Access Number )
13.7.4 | Support Import Filter
13.7.5 | Support DICOM Modality Worklist Standard
13.8 | Overlay Display on image
13.8.1 | Projection description
13.8.2 | Hospital or Institution Name, Date and Time of examination
13.8.3 | Radiographer Name
13.8.4 | Patient Name, Sex and Age
13.8.5 | Machine Name
13.8.6 | kV / mA / Time / mAs
13.8.7 | Feed-back mAs / Feed-back Time for AEC
13.8.8 | EI(Exposure Index)
14 | THE FOLLOWING MUST BE PROVIDED WITH THE SYSTEM INTERFACES AND CONNECTIVITY
14.1 | A "Conformance Statement" for all DICOM interfaces must be included. (the document must be attached)
14.2 | Interface DICOM: "Basic Print"
14.3 | Interface DICOM: "Modality Work-list"
14.4 | Interface DICOM: "Send as Service Class User (SCU)"
14.5 | Interface DICOM: "Receive as Service Class User (SCU)"
14.6 | Interface DICOM: "Retrieve as Service Class User (SCU)"
14.7 | Interface for checking the system status or for loading new
14.8 | Software to all components should be DICOM Interfaced.
14.9 | The modality shall include the DICOM Service Class license Image Storage: SCU and SCP
14.10. | The modality shall include the DICOM Service Class license Query/Retrieve SCU
14.11 | The modality shall include the DICOM Service Class license Print SCU
14.12 | The modality shall include the DICOM Service Class license Storage Commitment SCP
14.13 | The vendor shall include annexures of Proof of DICOM Conformance Statement
14.14 | The succesful bidder must ensure that the configuration between the processed images on the Digital Imaging Workstation and that transferred to the receivers of these images, such as PACS , Laser Printer , CD/DVD are properly calibrated/ reconfigured during installation, service and maintenance,repairs, and future software updates/upgrades.
15 | GENERAL
15.1 | Remote diagnostics feature for fault finding / troubleshooting must be available
15.2 | All computer systems supplied must have UPS protection with at least 20 minutes backup
15.3 | The Unit must link to existing laser printer
15.4 | The Unit must link to existing CD Robot
15.5 | The Unit offered must be able to connect a PACS/ RIS and HIS
15.6 | The Unit must be able to connect to any vendor
17 | RADIATION PROTECTION
17.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
17.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
17.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
17.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
17.5 | A radiation report must be automatically generated and transferable via DICOM.
17.6 | Should be FDA or CE approved product
17.7 | Electrical safety conforms to standards for electrical safety IEC-60601
17.8 | The unit must comply to IEC-61267
17.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
17.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
17.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
17.12 | Lead Glass window to be installed as per room requirement
17.13 | Safety aspects of Radiation dosage leakage should be spelt out
17.14 | DB boards and electrical filters to be installed
17.15 | QA TOOLS
17.16 | Display Exposure index values
17.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
17.18 | Patient Conventional face shield to be included.
17.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
17.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
17.21 | Interlocks to stop system movement when compression is applied must be provided.
17.22 | Interlock for premature release of the X-ray switch must be available
17.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
17.24 | Phantom for calibration should be provided.
18 | POWER SUPPLY
18.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
18.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
18.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
18.4 | A 3 phase in / out power supply must be offered. State details.
18.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
18.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
18.7 | The lifetime of UPS must be at least 15 years.
18.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
18.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
18.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
18.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
18.12 | Maintenance free batteries must be offered.
19 | AIR CONDITIONER
19.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
19.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
19.3 | The air conditioner must also be included in the maintenance plan.
19.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
19.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
19.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
19.7 | Air-conditioning must be included in Maintenance plan