## Sheet1
DRY LASER IMAGER
Item | TECHNICAL SPECIFICATIONS
1 | A good quality dry laser imager must be included in this offer. State the dimensions of the system offered.
2 | The unit must use either laser or thermal technology. No units using chemical technology will be considered.
3 | The printer must automatically perform a calibration test at start-up or on the request of the operator.
4 | The printer must include a self diagnostic capability which shall indicate at least the following error and system status conditions; film low, film empty, memory full, film feed error, printing, and alarm.
5 | The unit must allow daylight loading of film. State the loading capacity of the printer supply magazine.
6 | Multi-format printing on all film sizes must be possible. State the printing formats available on the offered printer.
7 | The printer must print the Department name, patient identification, date (dd/mm/yy) and time onto the film.
8 | Requests for printing must not compromise any additional workstation operation or performance.
Must include an online densitometer for constant image quality assessment including self calibration and optical density consistency as per SAHPRA QC test requirements
9 | The printer must be DICOM 3.0 compatible and enabled.
10 | It must be possible to print the following film sizes: 35cm X 43cm, 24cm X 30cm or 18cm X 24cm. The printer should be capable of supporting 2 film sizes on - line permanently.
11 | The throughput must be at least 120 (35cm x 43cm) films per hour. State the throughput of the dry laser imager offered. (ie number of 35cm x 43cm films per hour).
12 | It must be possible to send a print command from both the CR reader system and all optional workstations.
13 | POWER SUPPLY
13.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
13.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
13.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
13.4 | A 3 phase in / out power supply must be offered. State details.
13.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
13.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
13.7 | The lifetime of UPS must be at least 15 years.
13.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
13.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
13.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
13.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
13.12 | Maintenance free batteries must be offered.
14 | AIR CONDITIONER
14.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
14.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
14.3 | The air conditioner must also be included in the maintenance plan.
14.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
14.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
14.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
14.7 | Air-conditioning must be included in Maintenance plan