## Sheet1
FLUOROSCOPIC SYSTEM
ITEM | TECHNICAL SPECIFICATIONS
1. | X-Ray Generator
1.1 | Micro-processor controlled high frequency generator | #
1.2 | Nominal rating: > 80kW | #
1.3 | Shortest exposure time: 1msec | #
1.4 | Maximum tube current: ± 1000 mA | #
1.5 | Fluoroscopic programmed radiography must be possible | #
1.6 | Auto-setting of all parameters of at least 10 different examinations should be possible | #
1.7 | Anatomical programming must be included. State # of programmes | #
1.8 | Pulsed fluoroscopy for dose reduction should be available
1.9 | Automatic mains compensation | #
1.10 | Continuous self-check of all generator functions
1.11 | An integrated operator's console is required. | #
1.12 | Can the measuring fields be graphically displayed over the LIH image?
2. | X-Ray Tube:
2.1 | A high speed X-Ray tube is essential. State speed | #
2.2 | Dual focus with ratings compatible with the offered generator | #
State focal sizes.
2.2.1 | Large focus to be ± 1.0 mm | #
2.2.2 | Small focus to be ± 0.6 mm | #
2.3 | Anode material to be tungsten rhenium molybdenum compound with a maximum heat storage capacity of at least 1 500 kHU
2.4 | State anode heat dissipation HU/min | #
2.5 | Short term rating to be at least 45 kW and 85 kW | #
2.6 | Set of high tension cables to suit the installation | #
2.7 | The table must be able to perform tomography. Please state the angles offered and the speeds related to each angle.
3 | Integrated Table and C-Arm
3.1 | Remote controlled table/detector with integrated C-arm | #
3.1.1 | Table angulation: + 90˚ and -45˚ | #
3.2 | Table top must be X-ray translucent | #
3.3 | Table top length approximately 220cm x 55cm | #
3.4 | Range of table top/detector movements: | #
3.4.1 | Transverse at least 36 cm | #
3.4.2 | Longitudinal patient coverage of approx.195cm should be possible | #
3.5 | SID should be variable to position the detector close to the patient without tabletop or C-arm elevation | #
3.6 | Table height at lowest level should be ± 70 cm | #
3.7 | AP and PA imaging must be possible | #
3.8 | Programming for at least 50 preset positions in both PA and AP must be available | #
3.9 | C-arm angulation of approx. 35° cranial and 35°caudal | #
3.10 | C-arm angulation speed variable, at least 15 ˚/sec | #
3.11 | At least 90° RAO and 45° LAO must be possible | #
3.12 | An automatic collimator system | #
3.13 | The remote control panel must be fully integrated and include controls for generator, image acquisition and tabletop/C-arm | #
3.14 | All relevant generator and system information must be presented as clearly legible displays | #
3.15 | Controls for table top/C-arm movement should be available at the table side and the remote panel | #
3.16 | Collision sensors for patient safety must be incorporated within the system | #
3.17 | Maximum weight limit for the table should be at least 165kg | #
3.18 | The following table accessories are to be included:- | #
3.18.1 | Shoulder rests
3.18.2 | Mattress
3.18.3 | Hand grips
3.18.4 | Foot rest
3.18.5 | Compression cone
3.18.6 | Compensation filter pad for peripheral studies
4 | Digital Flat Panel Detectors
4.1 | A large field digital flat panel with a maximum input field of 43 x 43 cm is required. | #
4.1.1 | Direct conversion method will be an advantage. Tenderer to state and motivate what technology/ method is used | #
4.2 | Detector matrix must be at least 2048 x 2048 or greater, 16bit. State and motivate | #
4.3 | Processor matrix must be at least 2048 x 2048 or greater, 16 bit. State and motivate | #
4.4 | State acquisition field sizes in sq.cm | #
4.5 | Resolution must not be less than 2.5 lp/mm. State | #
4.6 | State and motivate pixel size. Should not be more than 180µm | #
4.7 | State DQE – Acquisition dose (%). Preference will be given to the lowest dose system | #
4.8 | State DQE – Fluoroscopy dose (%). Preference will be given to the lowest dose system | #
5 | T.V. Monitors
5.1 | Two high resolution TFT LCD monitors of at least 48 cm to be ceiling mounted in the examination room | #
5.2 | One at least 48 cm high line-rate TFT LCD monitor to be supplied at the control desk | #
5.3 | Automatic brightness control to be included | #
5.4 | State monitor output brightness: (cd/m2)
5.4.1 | Monitors in exam room | #
5.4.2 | Monitor at control panel | #
6 | DIGITAL IMAGE ACQUISITION AND PROCESSING SYSTEM
6.1 | A digital imaging system suitable for general fluorography
6.2 | Digital acquisition matrix:
6.2.1 | Digital radiography: Minimum of 2048 X 2048 (2 Kilo Matrix), 12 bit, 0.5 to 7.5 frames/sec
6.2.2 | Digital fluoroscopy should be possible at 1024X1024, 30 FPS
6.3 | At least 3 pulsed fluoroscopy modes must be available
6.4 | Single shot exposures must be possible
6.5 | Last image hold function and instant transfer of fluoroscopy image, both single image and dynamic, to digital storage on a hard disk with at least 4 GB capacity. State capacity
6.6 | Storage capacity of 20 000 frames in 1024 X 1024 matrix for immediate access must be available. State capacity
6.7 | Matrix on CD-R must be possible. State
6.8 | DICOM 3 conformance statement to be included DICOM Print must be possible
6.9 | Backup recording on CD-R should also be possible
6.1 | Connection for sending images to a laser imager should utilize a digital or DICOM printing environment
6.11 | Functionality must also include:-
6.11.1 | Measuring of distance, angle and area
6.11.2 | Gray scale inversion
6.11.3 | Window and centre control
6.11.4 | Text annotation
6.11.5 | Up to X2 zoom facility required complete with scroll, zoom and cine loop display
6.11.6 | Horizontal (R/L) and vertical (up/down) image flip
6.11.7 | Gamma curve selection
6.11.8 | 16 & 4 on 1 multi-frame display with study/series overview Split display for reference image comparison on 1 monitor
7 | IT INTEGRATION
7.1 | The system must have the following DICOM compatibility:
7.1.1 | DICOM print class
7.1.2 | DICOM storage class
7.1.3 | DICOM query/retrieve
7.1.4 | DICOM worklist
7.1.5 | The system must comply to the following IHE profiles (Integrating the Healthcare Enterprise)
7.1.6 | Scheduled Workflow (SWF)
7.1.7 | Patient Information Reconciliation (PIR)
8 | RADIATION PROTECTION
8.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
8.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded). | #
8.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems. | #
8.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions. | #
8.5 | A radiation report must be automatically generated and transferable via DICOM. | #
8.6 | Should be FDA or CE approved product | #
8.7 | Electrical safety conforms to standards for electrical safety IEC-60601
8.8 | The unit must comply to IEC-61267 | #
8.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa | #
8.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
8.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
8.12 | Lead Glass window to be installed as per room requirement
8.13 | Safety aspects of Radiation dosage leakage should be spelt out
8.14 | DB boards and electrical filters to be installed
8.15 | QA TOOLS
8.16 | Display Exposure index values
8.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
8.18 | Patient Conventional face shield to be included.
8.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
8.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
8.21 | Interlocks to stop system movement when compression is applied must be provided.
8.22 | Interlock for premature release of the X-ray switch must be available
8.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
8.24 | Phantom for calibration should be provided.
9 | POWER SUPPLY
9.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
9.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
9.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
9.4 | A 3 phase in / out power supply must be offered. State details.
9.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
9.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
9.7 | The lifetime of UPS must be at least 15 years.
9.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
9.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
9.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
9.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
9.12 | Maintenance free batteries must be offered.
10 | AIR CONDITIONER
10.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
10.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
10.3 | The air conditioner must also be included in the maintenance plan.
10.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
10.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
10.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
10.7 | Air-conditioning must be included in Maintenance plan