## Sheet1
ULTRASOUND MACHINE, PORTABLE, HIGH END AND COMPACT
ITEM | TECHNICAL SPECIFICATIONS | Weight
1 | General
1.1 | Unit should offer at least a 17 to 19 inch LCD high resolution colour monitor with integrated keyboard (laptop type). State the size and technology. | 2
1.2 | Indicate if system has a built-in UPS | 2
1.3 | The unit must come with a trolley
2 | The machine must be configurable to perform the following ultrasound investigations:
2.1 | Adult cardiac and paediatric | #
2.2 | Adult and paediatric transesophageal | 2
2.3 | Stress echo | #
2.4 | Abdominal | #
2.5 | Paediatric | #
2.6 | Vascular | #
2.7 | Obstretrical | #
2.8 | Gynaecological | #
2.9 | Small parts | #
2.1 | Breast | #
2.11 | Musculoskeletal | #
2.12 | Emergency medicine | 2
2.13 | Regional anaesthesia | 2
2.14 | Intergrated standby battery
3 | imaging modes
3.1 | • 2D | #
3.2 | • M-mode | #
3.3 | • Anatomical M-mode | #
3.4 | • Color M-mode | #
3.5 | • Color Power doppler imaging | #
3.6 | • Pulsed wave (PW) Doppler | #
3.7 | • High Pulse Repetition Frequency (HPRF) PW Doppler | #
3.8 | • Continuous wave (CW) Doppler | #
3.9 | • Freehand 3D imaging (available as an option) | 1
3.1 | •advanced quantification/ measurement software | 2
3.11 | • Invert and color invert | #
3.12 | • Color compare mode | #
3.13 | • Dual mode | #
3.14 | • Duplex for simultaneous 2D and Doppler | #
3.15 | • Triplex for simultaneous 2D, Doppler, and color or color power doppler | #
3.16 | • 2D and flow optimization signal processing | #
3.17 | • Tissue harmonic imaging (THI) | #
3.18 | • Reconstructed zoom with pan | #
3.19 | • High definition write zoom | #
3.20. | • Trapezoidal imaging | #
3.21 | • Contrast imaging (as an option) | 1
3.22 | • Adaptive Doppler | #
3.23 | • Adaptive color Doppler | #
3.24 | • Color tissue Doppler imaging | #
3.25 | • Pulsed wave tissue Doppler imaging | #
4 | 2D grayscale
4.1 | • Time Gain Compensation (TGC): pre-defined TGC curves optimized for consistently | 2
4.2 | • User adjustable Lateral Gain Control (LGC) | #
4.3 | • High definition zoom concentrates all image processing power into a user-defined area of interest | #
4.4 | • ability to combine high definition zoom with pan zoom | #
4.5 | • Cineloop image review | #
4.6 | • Selectable 2D compression settings | #
4.7 | • Sector size and steering control for sector and curved array | #
5 | image formats
5.1 | • Dual imaging with either independent cineloop buffers or split screen imaging | #
5.2 | • Live compare | #
5.3 | • 256 Gray scale levels or more | #
5.4 | • 2D acquisition frame rate of 400 frames/sec up to 800 frames/sec. State details. | 2
6 | M-mode
6.1 | • Available on all imaging transducers | #
6.2 | • Selectable sweeping rates | #
6.3 | • Time markers: minimum of 0.2 seconds up to a maximum of 2.0 seconds | 2
6.4 | • Acquisition zoom capability | #
6.5 | • Selectable display format prospective and retrospective | #
6.6 | • Cineloop review for retrospective analysis | #
7 | Anatomical M-mode
7.1 | • Available for all imaging transducers | #
7.2 | • Uses 2D image as a basis for M-mode analysis at a defined line, independent of transducer orientation | #
7.3 | • Anatomical M-mode trace can be generated or modified post freeze | #
8 | Tissue Doppler imaging
8.1 | • Allows high frame rate acquisition of tissue motion | #
8.2 | • Color gain, TGC, and LGC | 2
8.3 | • 15 to 20 color maps | 2
8.4 | • Tissue velocity (6cm to 50cm/s) | 2
9 | Pulsed wave Doppler
9.1 | • Available on all imaging transducers | #
9.2 | • Adjustable sample volume size:(e.g. 0.5-20 mm) -state size | 2
9.3 | • Triplex / real time duplex mode of operation | #
9.4 | • Simultaneous 2D, color Doppler or CPA, pulsed Doppler | #
9.5 | • automatic optimization button / knob - adjusts scale, baseline and doppler gain | 2
10 | Continuous wave Doppler
10.1 | • Steerable through 90° sector | 2
10.2 | Maximum velocity range: 15 sec (transducer dependent)
11 | Spectral Doppler
11.1 | • Display annotations including: | #
11.1.1 | Doppler mode
11.1.2 | scale (cm/sec)
11.1.3 | Nyquist limit
11.1.4 | wall filter setting
11.1.5 | gain
11.1.6 | acoustic output status
11.1.7 | sample volume size
11.2 | • Angle correction with automatic velocity scale adjustment | #
11.3 | • Adjustable velocity display ranges | #
11.4 | • Normal/invert display around horizontal zero line | #
11.5 | • Selectable sweep speeds | #
11.6 | • Selectable low-frequency signal filtering with adjustable wall filter settings | #
11.7 | • Selectable grayscale curve for optimal display | #
11.8 | • Selectable display format of prospective and retrospective | #
11.9 | • 256 Gray scale levels or more | #
11.10. | • Post-processing in Pulsed Wave (PW) frozen mode includes map, baseline,sweep speed, invert, compress, reject, and chroma | #
12 | Color Doppler
12.1 | • Adaptive mode adjusts Doppler frequency and sensitivity based on color ROI placement within image available on all imaging transducers | #
12.2 | • Cineloop review with full playback control | #
12.3 | • Advanced motion suppression with intelligent algorithms; to eliminate color motion artifacts. State details. | 2
12.4 | • Trackball-controlled color region of interest: size and position | #
12.5 | • Maps, filters, color sensitivity, line density, smoothing, echo, color persistence, gain and baseline optimized automatically by exam type or is user selectable | #
12.6 | • Velocity and variance displays | #
12.7 | • Color invert in live and frozen imaging | #
12.8 | • User selectable smoothing control | #
12.9 | • User selectable persistence control | #
12.1 | • Color/2D line density control | #
13 | Tissue harmonic imaging
13.1 | • Available on all imaging transducers | #
13.2 | Extends high performance imaging capabilities to all patient body types
14 | Color Power doppler
14.1 | Directional power flow available | #
14.2 | • Available on all imaging transducers | #
14.3 | • Cineloop review | #
14.4 | • Multiple color maps | #
14.5 | • Individual controls for gain, Pulse Repetition Frequency (PRF), baseline, filters, sensitivity, | #
14.6 | • Dynamic motion differentiation | #
14.7 | • Adjustable region of interest: size and position | #
14.8 | • User selectable Persistence button / knob. | 2
14.9 | • User selectable Blending button / knob. | 2
14.10. | Highly sensetive mode for small vessel visualization
15 | Connectivity
15.1 | • One USB ports on control panel | 2
15.2 | • At least 80 GB hard drive space or more | 2
15.3 | • Internal slot-load CD/DVD RW drive | 2
15.4 | • DICOM print, store, and storage commitment | #
15.5 | • DICOM structured reporting capability | 2
15.6 | • Modality worklist | #
15.7 | • A DICOM images must exported with the DICOM reader. | 2
15.8 | The DICOM conformance statement must be included. | 2
15.9 | • Export data as PC compatible or DICOM files | #
15.1 | • Ethernet at 100 Mb/second | 2
15.11 | • B/W printers to be supplied. | 1
15.12 | • The ultrasound unit must be DICOM and HL7 compliant and enabled to be PACS ready without any additional costs. | 2
16 | Probes: the following probes must be included in the price. Give prices for individual probes.
16.1 | The selected frequency of the imaging transducer must be displayed on the screen. | #
16.2 | Three transducers must be included and the end user must be able to select the final transducers Clearly State all available transducers with prices per transducers (swapping of products/transducers and software to be allowed with equitable costs). Details are stated | 2
16.3 | The transducers selected by the end-user must be supplied with the relevant software. | 1
16.4 | The relevant and compatible software are the following in line with the transducers selected: MSK, TEE adult, TEE pediatric, pediatric TTE, adult TTE, small parts, vascular, breast, abdominal and obstretrics & gynaecology. | 2
16.5 | additional probes and software must be available as optional items. Optional items must be listed and priced separately. | 1
16.6 | Hockey stick optional | 1
16.7 | Convex Probe (Curvilinear). The MHz range are stated (within 1-6 MHz) | #
16.8 | Microconvex for neonatal/ paediatrics.The MHz range are stated (4-8) | 2
16.9 | Microconvex Transvaginal/transrectal. The MHz range are stated (within 3-8 MHz) | 2
16.1 | Linear probe (vascular) The MHz range are stated (4-12) | 2
16.11 | Linear probe (MSK)The lower MHz range must be 5-6 MHz and the upper range must be between 14 and 21 MHz. | 2
16.12 | Adult Cardiac Probe. State the MHz range (1-4MHz) | 2
16.13 | Paediatric Cardiac Probe. State the MHz range (2-10 MHz) | 2
16.14 | TEE Probe, adult. The lower MHz range must be 2MHz and the upper range must be 9MHz. | 2
16.15 | TEE Probe, pediatric. The lower MHz range must be 3MHz and the upper range must be 11MHz. | 2
17 | SOFTWARE
17.1 | Compound Imaging ( Cross Beam angles) | #
17.2 | Speckle Reduction (Minimum of 4-5 steps ). Refines images(reduces noise) | #
17.3 | System must have software to enhance needle visualization during biopsy | 2
17.4 | Measurement and Report software packages must be configurable | 1
17.5 | Details of tools/software provided as a standard are stated | 1
17.6 | The standard software packages supplied must be fully upgradable . | #
17.7 | The cost of updating and upgrading all standard software packages supplied , during warranty and maintenance period must be included in the offer . It must be clearly outlined as a line item in the offer documentation. | #
17.8 | A rechargeable battery with a charger must be supplied. | 1
17.9 | Fully charged battery must have an actual time use of one hour or more. | 2
17.10. | The weight of the portable ultrasound unit must be less than five kg. | 2
17.11 | A multi port (for transducer connectivity) cart with castors with two wheel lock brakes | 2
17.12 | A compact carry case with wheels for the ultrasound unit must be supplied. | 2
17.13 | Paediatric and adult ECG cables (three lead) must be supplied. | 2
18 | OPTIONAL ITEMS
18.1 | USB PORT additional | 1
18.2 | Any additional tools/software not included in the standard package must be clearly listed and priced. | 1
18.3 | Elastography software | 1
18.4 | 3D software | 1
18.5 | 4D software | 1
19 | POWER SUPPLY:
19.1 | The equipment offered shall operate from a standard 220 to 240 Volt A.C. 50Hz mains supply. The bidder must ensure that the product being quoted for is fitted with a 15 Amp sealed plugtop S.A.N.S approved
19.2 | The unit must be supplied with an on-line uninterrupted power supply (U.P.S.) to provide minimum of 50 to 60 minutes mains backup power.
19.3 | The UPS should not adversely affect the image quality.
19.4 | This unit must be included in the bid price and be covered by the guarantee and maintenance plan.
19.5 | The UPS should not adversely affect the image quality.