## Sheet1
ULTRASOUND MACHINE, PORTABLE, SMALL AND COMPACT, HIGH END
ITEM | TECHNICAL SPECIFICATIONS | Weight
1 | The machine must be configurable to perform the following ultrasound investigations:
1.1 | Adult cardiac and paeditric | #
1.2 | Adult transesophageal | 2
1.3 | Stress echo | #
1.4 | Abdominal | #
1.5 | Pediatric | #
1.6 | Vascular | #
1.7 | Obstetrical | #
1.8 | Gynecological | #
1.9 | Small parts | #
1.10 | Breast | #
1.11 | Musculoskeletal | #
1.12 | Emergency medicine | #
1.13 | Regional anesthesia | 2
1.14 | Integrated Standby Battery | #
2 | imaging modes
2.1 | • 2D | #
2.2 | • M-mode | #
2.3 | • Anatomical M-mode | #
2.4 | • Color M-mode | #
2.5 | • Color Power doppler imaging | #
2.6 | • Pulsed wave (PW) Doppler | #
2.7 | • HPRF PW Doppler | #
2.8 | • Continuous wave (CW) Doppler | 2
2.9 | • Freehand 3D imaging (available as an option) | 1
2.10 | •advanced quantification/ measurement software | 2
2.11 | • Invert and color invert | #
2.12 | • Color compare mode | #
2.13 | • Dual mode | #
2.14 | • Duplex for simultaneous 2D and Doppler | #
2.15 | • Triplex for simultaneous 2D, Doppler, and color or color power doppler | #
2.16 | • 2D and flow optimization signal processing | #
2.17 | • Tissue harmonic imaging (THI) | #
2.18 | • Reconstructed zoom with pan | #
2.19 | • High definition write zoom | #
2.20 | • Trapezoidal imaging | #
2.21 | • Contrast imaging (as an option) | 1
2.22 | • Adaptive Doppler | #
2.23 | • Adaptive color Doppler | #
2.24 | • Color tissue Doppler imaging | #
2.25 | • Pulsed wave tissue Doppler imaging | #
3 | 2D grayscale
3.1 | • Smart TGC: pre-defined TGC curves optimized for consistently | #
3.2 | • User adjustable LGC control | #
3.3 | • High definition zoom concentrates all image processing power into a user-defined area of interest | #
3.4 | • possible to combine high definition zoom with pan zoom | #
3.5 | • Cineloop image review | #
3.6 | • Selectable 2D compression settings | #
3.7 | • Sector size and steering control for sector and curved array | #
4 | image formats
4.1 | • Dual imaging with either independent cineloop buffers or split screen imaging | #
4.2 | • Live compare | #
4.3 | • 256 (10-12 bits) discrete gray levels or better | #
4.4 | • state 2D acquisition frame rate (preferably 400 up to 755 frames/sec) | #
5 | M-mode
5.1 | • Available on cardiac imaging transducers | #
5.2 | • Selectable sweeping rates | #
5.3 | • Time markers: 0.2 and 1.0 seconds | 2
5.4 | • Acquisition zoom capability | #
5.5 | • Selectable display format prospective or retrospective | #
5.6 | • Cineloop review for retrospective analysis | #
6 | Anatomical M-mode
6.1 | • Available for cardiac imaging transducers | #
6.2 | • Uses 2D image as a basis for M-mode analysis at a defined line, independent of transducer orientation | #
6.3 | • Anatomical M-mode trace can be generated or modified post Freeze | #
7 | Tissue Doppler imaging
7.1 | • Allows high frame rate acquisition of tissue motion | #
7.2 | • Color gain, TGC, and LGC | 2
7.3 | • 15 to 20 color maps | #
7.4 | • Velocity (6cm to 50cmcm/s) | #
8 | Pulsed wave Doppler
8.1 | • Available on all imaging transducers | #
8.2 | • Adjustable sample volume size:(e.g. 0.5-20 mm) -state size | #
8.3 | • Simultaneous or duplex mode of operation | #
8.4 | • Simultaneous 2D, color Doppler or CPA, pulsed Doppler | #
8.5 | • automatic optimization - adjusts scale, baseline and doppler gain | #
9 | Continuous wave Doppler
9.1 | • Steerable through 90° sector | 2
9.2 | • Maximum velocity range: 15sec (transducer dependent) | #
10 | Spectral Doppler
10.1 | • Display annotations including Doppler mode, scale (cm/sec),Nyquist limit, wall filter setting, gain, acoustic output status, sample volume size, normal/inverted, angle correction,grayscale curve | #
10.2 | • Angle correction with automatic velocity scale adjustment | #
10.3 | • Adjustable velocity display ranges | #
10.4 | • Normal/invert display around horizontal zero line | #
10.5 | • Selectable sweep speeds | #
10.6 | • Selectable low-frequency signal filtering with adjustable wall filter settings | #
10.7 | • Selectable grayscale curve for optimal display | #
10.8 | • Selectable display format prospective or retrospective - state options | #
10.9
10.10 | • 256 (10 t0 12 bits) discrete gray levels | #
10.11 | • Post-processing in PW frozen mode includes map, baseline,sweep speed, invert, compress, reject, and Chroma | #
11 | Color Doppler
11.1 | • Adaptive mode adjusts Doppler frequency and sensitivity based on color ROI placement within image available on all imaging transducers | #
11.2 | • Cineloop review with full playback control | #
11.3 | • Advanced motion suppression with intelligent algorithms; to eliminate color motion artifacts. Give details | 2
11.4 | • Trackball-controlled color region of interest: size and position | #
11.5 | • Maps, filters, color sensitivity, line density, smoothing, echo, write priority, color persistence, gain and baseline optimized automatically by exam type or is user selectable | #
11.6 | • Velocity and variance displays | #
11.7 | • Color invert in live and frozen imaging | #
11.8 | • User selectable smoothing control | #
11.9 | • User selectable persistence control | #
11.1 | • Color/2D line density control | #
12 | Tissue harmonic imaging
12.1 | • Available on all imaging transducers | #
12.2 | • Extends high performance imaging capabilities to all patient body types | #
13 | Color Power doppler
13.1 | Directional power flow available | #
13.2 | • Highly sensitive mode for small vessel visualization | #
13.3 | • Available on all imaging transducers | #
13.4 | • Cineloop review | #
13.5 | • Multiple color maps | #
13.6 | • Individual controls for gain, PRF, baseline, filters, sensitivity, | #
13.7 | • Dynamic motion differentiation | #
13.8 | • Adjustable region of interest: size and position | #
13.9 | • User-selectable persistence | #
13.10 | • User-selectable blending | #
14 | Connectivity
14.1 | • Two USB ports on control panel | 2
14.2 | • At least 80 GB internal hard drive space or more | #
14.3 | • Internal slot-load CD/DVD RW drive | 2
14.4 | • DICOM print, store, and storage commitment | #
14.5 | • DICOM structured reporting for cardiac and obstetrics | 2
14.6 | • Modality worklist | #
14.7 | • DICOM reader saved onto media | #
14.8 | • Export data as PC-compatible or DICOM files | #
14.9 | • Ethernet at 100 Mb/second | 2
14.1 | • USB to serial converter adapter | 1
14.11 | • Support for optional small B/W and color printers | 1
14.12 | • Must be PACS ready with no added cost to connect to PACS | #
15 | Probes: the following probes must be included in the price. Give prices for individual probes.
15.1 | Adult TEE (offer as optional) | 1
15.2 | Vascular linear (L12-3) | #
15.3 | Abdominal curviliar (C5-1) | #
15.4 | Hockey stick (L15-7) optional | 1
15.5 | Microconvex/ pediatric abdomen (C8-5) | #
15.6 | Sector transducer 1.5Mhz to 4.5 Mhz (adult cardiac) | 2
15.7 | Sector transducer 3 to 8Mhz for paediric cardiac | 2
16 | Power Supply
16.1 | The unit must be supplied with an on-line uninterrupted power supply (U.P.S.) to provide minimum ten 10 to 20 minutes mains backup power.
16.2 | The UPS should not adversely affect the image quality.
16.3 | This unit must be included in the bid price and be covered by the guarantee and maintenance plan.