## Sheet1
ULTRASOUND UNIT (HIGH END)
TECHNICAL SPECIFICATION
ITEM | IMAGING MODALITIES
1.1 | This unit must be a fully digital high end Colour Doppler unit with pulsed wave Doppler and Continuous wave Doppler.
1.2 | The unit must have a standby mode for quick start up & shutdown.
1.3 | The unit must have 3 active probe connectors with electronic switching.
1.4 | The unit must be ergonomically designed with an elevating adjusting panel to reduce repetitive stress injuries.
1.5 | The unit must offer a 48 cm or better high resolution flat panel colour monitor which must be able to move up and down and from side-to-side.
1.6 | A cine loop function of at least 30sec must be offered for B-mode and D - mode.
1.7 | The unit must be fully DICOM/PACS compatible. DICOM verification, worklist, structured reporting, storage and print must be available as standard.
1.8 | The unit offered must have at least 29,000 channels.
1.9 | The unit must offer at least 512 line density in B-mode.
1.1 | Phase Inversion Tissue Harmonics must be offered on all probes
1.11 | Spatial compounding offering multiple lines of sight must be offered to improve contrast resolution and speckle reduction.
1.12 | The system offered must have at least a 500GB hard drive.
1.13 | The unit must offer TGC in not less than 8 steps.
1.14 | Pan-zoom enlargement must be at least 10 times read & write. Image quality must not be compromised by the zoom function.
1.15 | Easy programmable image processing must be in not less than 8 steps to enable the user to change dynamic range, persistence, contrast, enhancement, smoothing and post-processing.
1.16 | Independent M-mode gain adjustment must be available.
1.17 | B-mode frame rate of at least 498 fps
1.18 | Colour frame rate higher than 200 fps
1.19 | B-Mode image on the linear probe must be electronically steerable.
1.20. | Trapezoidal imaging must be available on linear probes.
1.21 | Sector FOV must be at least 90 deg.
1.22 | Maximum depth must not be less than 28cm.
1.23 | System dynamic range must be greater than 160dB.
1.24 | One button B-mode image optimisation must be possible.
1.25 | Panoramic software must be available in B-mode & Colour.
1.26 | Real-time 4D imaging of at least 30 volumes per second. Trans-abdominal and trans-vaginal 4D transducers and software to be included.
1.27 | Images of video clips and still images must be able to be stored on hard drive. Measurements on stored images must be possible.
1.28 | Loop / Clip storage of up to 1 minute must be possible.
1.29 | Please state any other additional features/technologies that are available on the unit offered.
2 | PULSED WAVE (PW) & CONTINUOUS WAVE DOPPLER (CW).
2.1 | PW with a one button Doppler optimisation must be possible.
2.2 | PW Doppler on all imaging probes.
2.3 | The max detectable velocity for PW must be stated.
2.4 | Maximum PRF must be equal to or greater than 40kHz.
2.5 | Independent gain adjustments for PW & CW must be available.
2.6 | Independent angle steering up to 20° for colour and PW must be offered on Linear Probes.
2.7 | Colour pallet offering 7 types independent of B-Mode, Pulse Doppler & Continuous wave doppler pallet must be available.
2.8 | Sampling with Doppler gate must be varied between 1 – 20mm.
2.9 | Display of Doppler scan modes side by side and up and down must be included .
2.10. | Post-processing of the Doppler trace must be possible in freeze mode.
2.11 | The unit must offer real time triplex mode on PW.
3 | COLOUR DOPPLER
3.1 | At least 8 colour maps must be available for colour Doppler.
3.2 | Colour maps must be able to be selected after frozen mode.
3.3 | Independent colour gain adjustment must be +15dB to -15dB in 0.2dB steps.
3.4 | Maximum colour frame rate must not be less than 200 frames p/sec.
3.5 | Colour flow must operate to the full depth of 28cm.
3.6 | No deterioration of B-mode must be noticeable in colour mode.
3.7 | There must be a minimum of 7 processing selections i.e. wall filters, colour averaging etc. State the number of selections and list the possible selections.
3.8 | The user must be able to select colour maps and processing selections. It must be possible to store these preset selections.
3.9 | The colour box must move with the Doppler gate to reduce scanning time.
3.10. | Power Doppler must be available to display the existence of blood flow with high sensitivity.
4 | TRANSDUCERS
4.1 | All probes required must be offered and included in the total bid price.
4.2 | All imaging transducers must be multi-frequency or broadband.
4.3 | Tissue Harmonics should be available on all probes.
4.4 | The selected frequency of the imaging transducer must be displayed on the screen. There should be at least B-mode frequencies, THI frequencies, PW frequencies and colour frequencies.
4.5 | All transducers must be digitally controlled so that continuous beam forming is possible.
4.6 | It must be possible to connect 3 probes simultaneously.
4.7 | The following transducers must be offered and included in the total bid price.
4.7.1 | Multi-frequency convex trans-abdominal transducer with Tissue Harmonics for Abdominal, O&G and vascular applications: 2 - 5 MHz central frequency.
4.7.2 | Wideband Linear transducer 5- 14MHz for general imaging
4.7.3 | High density Linear probe 18 MHz for breast and superficial imaging
4.7.4 | 5 – 9 MHz intercavity transducer for rectal and vaginal scanning
4.7.5 | Adult cardiac probe 5 MHz and Pediatric cardiac probe 7.5 MHz
5 | MEASUREMENTS (add for cardiac)
5.1 | 16 callipers (8 channel) for measuring distance angle area volume must be included.
5.2 | The unit must offer storage facility for up to 2000 patient fetal growth measurement data. This information must be retrievable for follow-ups at a later stage.
5.3 | 50th percentile growth graphs must be offered on all measurements for example BPD, FL, AC, Fetal Weight, HC. Information saved must be retrievable for all follow-up at a later stage.
5.4 | Display of graphs must be available to compare measurements of twins, triplets, quads.
5.5 | Unit must offer user programmable annotation and measurement labels to reduce scanning time.
5.6 | Automatic trace measurements must be possible on real-time and frozen Doppler traces.
5.7 | Supply details of any other measurement capabilities that are available on the system offered.
6 | RECORDING DEVICES:
6.1 | A black and white thermal printer must be included in the offer.
6.2 | A digital colour thermal printer must be offered as an option.
6.3 | Internal CD/DVD read / write must be offered as standard. It must be possible to write images in DICOM and PC format when storing to CD/DVD/PACS.
6.4 | A DVD writer/recorder for storage/recording of real time loops and still images must be offered as an option.
6.6 | A USB port for image transfer to a memory stick must be available.
7 | POWER SUPPLY:
7.1 | The equipment offered shall operate from a standard 220 to 240 Volt A.C. 50Hz mains supply. The bidder must ensure that the product being quoted for is fitted with a 15 Amp sealed plugtop S.A.N.S approved
7.2 | The unit must be supplied with an on-line uninterrupted power supply (U.P.S.) to provide minimum of 50 to 60 minutes mains backup power.
7.3 | The UPS should not adversely affect the image quality.
7.4 | This unit must be included in the bid price and be covered by the guarantee and maintenance plan.
7.5 | The UPS should not adversely affect the image quality.