## Sheet1
ULTRASOUND UNIT (MID-RANGE)
ITEM | Tender Specification | Weight
1 | Imaging Modes
1.1 | Unit must be latest model with latest technology. Details are stated. | #
1.2 | Ultrasound applications for abdominal, obstetrics, gynecological, small parts, vascular / peripheral vascular, musculoskeletal, urological, transcranial, cardiac and elastography . | #
1.3 | Unit should be capable of offering a frame rate from 400 frames per second and above. | #
1.4 | The colour frame rate should be 400 or above. State maximum colour frame rate. | 2
1.5 | Unit should offer a 42 cm or better high definition LCD colour monitor. | #
1.6 | Ergonomic design keyboard. Details must be stated. | 2
1.7 | Brightness / contrast adjustment on monitor. | #
1.8 | Programmable panel with changeable labels must be offered | #
1.9 | Graphical user interphase (user friendly and exam specific) | #
1.10. | Height adjustable user interphase | #
1.11 | Unit must be housed on a ergonomic and compact mobile cart with lockable wheels, for easy maneuvaribility | #
1.13 | Post processing must be available on frozen images ( Gain, Gray Map,Chrom Color map, Baseline, Remove color) . Details are stated. | #
1.14 | The gray scale level for image display modes, should be 256 shades and above. | #
1.15 | Touch screen Control Panel 8 inch and above. | #
1.16 | Unit should offer emergency log on. | #
2 | Functionality
2.1 | 12 bit and above digital converter for dynamic image / live image (beam former). | #
2.2 | B -mode must have trapezoid function. | #
2.3 | B mode must have panoramic view function. | #
2.4 | Autocalibration for measurements on panoramic image must be standard. | #
2.5 | B mode with broad band harmonic for increased resolution from near and far field. | #
2.6 | B-Mode image on linear transducer to be electronically steerable. | #
2.7 | Image format to operate in dual and quad screen. | #
2.8 | Independent gain control for Pulsed Wave Doppler. | #
2.9 | Scanning Operating Modes: B-mode, M-mode, Pulsed Wave Doppler (Spectral), Continuous Wave Doppler, Power Doppler & Color Doppler, 2D. | #
2.1 | Pulse Repetition Frequency (PRF) / Velocity must be adjustable in increments. | #
2.11 | Pulse Repetition Frequency (PRF) / Velocity values must be displayed on the screen. | 2
2.12 | Reference Frequency (REF) must be displayed. | 2
2.13 | Wall filter must be adjustable. | 2
2.14 | The sampling gate must be adjustable: range between 0.3mm and 15mm. | #
2.15 | Directional Power Doppler | 2
2.16 | M-mode / Doppler review (cine playback or loop) of 10 seconds and above. Details are stated. | 2
2.17 | M-mode / Doppler to be steerable. | #
2.18 | Independent B-mode and Color Doppler gain adjustment must be available. | #
2.19 | Independent adjustment of each mode for image optimization of each mode during triplex mode . | #
2.2 | The Maximum scanning depth must be 30 cm and above. Details are stated. | #
2.21 | One button image optimization must be standard resolution | 2
2.22 | Internal Hard drive with minimum usable image storage space of 250GB or higher. You must be able to store videos and still images on internal hard drive. | 2
2.23 | Auto deletion of images should be configurable | #
2.24 | Notice warning be included when minimum image storage capacity is reached on the system. The actual minimum storage capacity in GB and total number of images is stated in supporting documents provided. | #
2.25 | Black and white digital thermal printer attached to the unit must be included | #
2.26 | an Active DVD Writer/Recorder attached to the unit must be included | 1
2.27 | 2x USB Port (External) should be an integral part the unit for storage/recording of real time loops (video ) and still images. 2x Internal USB for Printer & memory stick use. | 2
2.28 | Software to ensure functionality of DVD Writer/Recorder with all types of computers/laptops must be included as a standard. | 2
2.29 | The system must be DICOM 3 compliant. | 2
2.3 | Advanced DICOM/PACS including full worklist must be offered, and to be included in Main Unit. | 2
2.31 | Export of images from the system must be possible via LAN, USB, DVD, (including transfer of images to PACS) . The methods shall be stated in supporting documents. | #
2.32 | the following storage formats must be possible on unit: JPEG, Bitmap, Tiff, AVI, DICOM | #
2.33 | On screen measurements must be possible on real time examination. | #
2.34 | 50th percentile growth graphs must be offered on all measurements for example BPD, FL, AC, Fetal Weight, HC. This information must be retrievable for all follow-up visits, to detect abnormal growth of foetus early | #
2.35 | Display of graphs must be available to compare measurements of twins , triplets in single and quad display. | #
2.36 | User programmable exam preset capability | #
2.37 | Unit must offer user programmable annotation for quick labelling. | #
2.38 | Storage and review of patient data must be available for comparative exams. Details are stated on the options you offer.if there is functional PACS. | 1
3 | Transducers
3.1 | The selected frequency of the imaging transducer must be displayed on the screen. | #
3.2 | Three transducers must be included and the end user must be able to select the final transducers Clearly State all available transducers with prices per transducers (swapping of products/transducers and software to be allowed with equitable costs). Details are stated. | 2
3.3 | Additional probes and software must be available as optional items. Optional items must be listed and priced separately. | 2
3.4 | Types and frequency range of transducers:
3.4.1 | Convex Probe (Curvilinear). The MHz range are stated (within 1-6 MHz) | #
3.4.2 | Microconvex for neonatal / paediatrics.The MHz range are stated (4-8) | 2
3.4.3 | Microconvex Transvaginal / transrectal. The MHz range are stated (within 3-8 MHz) | #
3.4.4 | Linear probe (vascular) The MHz range are stated (4-12) | 2
3.4.5 | Linear probe (MSK)The lower MHz range must be 5-6 MHz and the upper range must be between 14 and 21 MHz. | 2
3.4.6 | Adult Cardiac Probe. State the MHz range (5 MHz). Must have pulse and continuous Wave Doppler software | 2
3.4.7 | Paediatric Cardiac Probe. State the MHz range (7.5-10 MHz). Must have pulse and continuous Wave Doppler software | 2
3.4.8 | Minimum 3 x standard Probe holders. | #
3.4.9 | Minimum 3 x Active probe ports. Details of total number of active probe ports are stated. | #
3.4.10 | Foot rest as a standard. | 1
3.4.11 | Chair, ergonomic, height adjustable with lockable castors for the Health professional. | 2
3.4.12 | Printer mounted on the machine. | 1
3.4.13 | Cable tray /hooks cable support system. | 1
3.4.14 | Gel warmer. | 2
3.4.15 | Easy removable air filters (by end users). | 2
3.4.16 | Ultrasound Quality Control Phantom (swapping of phantoms, products and software to be allowed with equitable costs). Details are stated. | 2
3.5 | SOFTWARE:
3.5.1 | Compound Imaging ( Cross Beam angles). | #
3.5.2 | Speckle Reduction (Minimum of 4-5 steps ). Refines images(reduces noise) | #
3.5.3 | System must have software to enhance needle visualization during biopsy | 2
3.5.4 | Measurement and Report software packages must be configurable | 1
3.5.5 | Details of tools/software provided as a standard are stated | 1
3.5.6 | The standard software packages supplied must be fully upgradable . | #
3.5.7 | The cost of updating and upgrading all standard software packages supplied , during warranty and maintenance period must be included in the offer . It must be clearly outlined as a line item in the offer documentation. | #
4 | OPTIONAL ITEMS
4.1 | Fold down monitor for easy transportation. | 1
4.2 | USB Port, additional | 1
4.3 | Any additional tools / software not included in the standard package must be clearly listed and priced. | 1
4.4 | Side storage bin | 1
4.5 | Elastography software | 1
4.6 | 3D software | 1
4.7 | 4D software | 1
4.8 | Hard Drive of 5TB for additional storage/archiving is included.
5 | POWER SUPPLY:
5.1 | The equipment offered shall operate from a standard 220 to 240 Volt A.C. 50Hz mains supply. The bidder must ensure that the product being quoted for is fitted with a 15 Amp sealed plugtop S.A.N.S approved
5.2 | The unit must be supplied with an on-line uninterrupted power supply (U.P.S.) to provide minimum of 50 to 60 minutes mains backup power.
5.3 | The UPS should not adversely affect the image quality.
5.4 | This unit must be included in the bid price and be covered by the guarantee and maintenance plan.
5.5 | The UPS should not adversely affect the image quality.