## Sheet1
ULTRASOUND MACHINE, HIGH END, ECHOCARDIOGRAM UNIT WITH 4D/3D CAPABILITY WITH TEE
ITEM | TECHNICAL SPECIFICATIONS | Weight
1 | ULTRASOUND HIGH END ECHOCARDIOGRAM UNIT WITH 4D\3D CAPABILITY
1.1 | Top of the range High-End Ultrasound with Echocardiography | #
1.2 | Digital broadband beam | #
1.3 | Digital channels more than 500 000. | 2
1.4 | 2D Imaging | #
1.5 | Real time 4D (Live 3D) using matrix transducer technology for instantaneous volume rendering of cardiac anatomy | #
1.6 | Live biplane/ Triplane Imaging using Matrix transducers (infinite plane adjustments) | #
1.7 | Dynamic range of 250dB and above | #
1.8 | Real-Time Compounding (Spatial Compound) | #
1.9 | Speckle Reduction Software | #
1.10. | Speckle Trace/ Tracking Software (for wall motion abnormalities) | #
1.11 | Adaptive Image Processing for noise and artifact reduction | #
1.12 | Ergonomic transducers designs with lightweight flexible cables | #
1.13 | Color Frame Rates of 650 fps and above | #
1.14 | 2D Frame rates of 800 fps and above | #
1.15 | High-resolution interactive graphical color touch panel with adjustment for various ambient light conditions | 2
1.16 | Fully articulating flicker-free at least 19-inch to 20-inch wide format high resolution flat panel TFT with nearly infinite positioning adjustments | 2
1.17 | Fully articulating control panel, including height, swivel, and slide | #
1.18 | Must have minimum of four active transducer /probe ports and one compulsary non- imaging transducer (pencil probe) port. Thus 5 ports . | #
1.19 | Min /Max velocity for PW Doppler | #
1.2 | Min /Max velocity for CW Doppler | #
1.21 | Min /Max PRF for PW Doppler | #
1.22 | Intelligent one-button optimization in 2D and Doppler modes | #
1.23 | Integrated Standby Battery | #
1.24 | Unit must be housed on a lightweight mobile cart with lockable wheels. | 2
2 | CLINICAL APPLICATION IN SOFTWARES
2.1 | The echocardiogram machine must have the following software applications (protocols):
2.1.1 | Adult Echocardiography | #
2.1.2 | Pediatric Echocardiography (inclusive of neonatal) | #
2.1.3 | Stress Echocardiography (Treadmill and Chemical) | #
2.1.4 | 2D Transesophageal Echocardiography | #
2.1.5 | 4D / Live 3D Transthoracic Echocardiography (TTE) | #
2.1.6 | 4D / Live 3D Transesophageal Echocardiography | #
2.2 | Full Cardiac Echocardiogram Software Package | #
2.3 | Full Vascular Software Package | #
2.4 | The standard software packages supplied must be fully upgradable . | #
2.5 | The cost of updating and upgrading all standard software packages supplied , during warranty and maintenance period must be included in the offer . It must be clearly outlined as a line item in the offer documentation. | #
2.6 | Automated Intima Thickness Measurement | 2
3 | IMAGING MODES AND IMAGE DISPLAY
3.1 | The operating modes should include: 2D, M, 2D/M, 4D / Live 3D, PWD, High PRF, Color Doppler, Power Doppler, Directional Power Doppler, steerable CWD, Triplex and Harmonic with pulse inversion technology. | #
3.2 | Tissue Doppler with frame rate of 600 fps and above. | 2
3.3 | M mode color Doppler and M mode tissue Doppler | #
3.4 | Innovative Transducer crystal technology with a pure wave (or similar) capability for greater acoustic efficiency and bandwidth | 2
3.5 | Innovative Transducer crystal technology with active matrix array capability for greater acoustic efficiency | 2
3.6 | 4D / Live 3D Volume Imaging using matrix transducer | #
3.7 | 4D / Live 3D Zoom | #
3.8 | 4D / 3D Full volume | #
3.9 | 4D / Live 3D Quantification with MPR and islice. | #
3.1 | 4D / Live 3D display of dynamic 3D rendering and LV true volumes. | #
3.11 | Quantification of the mitral valve using the 4D / live 3D transesophageal transducer. | #
3.12 | Live (on screen) automatic Doppler Analysis | 2
3.13 | Automatic Doppler Analysis on still images. (autotrace) | #
3.14 | Tissue Specific Imaging Application Software | #
3.15 | Onboard quantification for :
3.15.1 | a. Automated Strain Imaging and Speckle tracking. | #
3.15.2 | b. Display of automated borders | #
3.15.3 | c. Display of Mitral valve annular motion | #
3.17 | Dual Live display with BW reference image and color Doppler image | #
3.18 | Colorization of volume imaging | #
3.19 | Trapezoid scanning providing convex view on linear probes. | #
3.2 | High Definition Write Zoom and Read Zoom with pan feature | #
3.21 | Scanning depth ranges for adult echocardiac probes of a minimum depth between 0 cm- 2cm, and a maximum depth between 30 cm-40cm | #
3.22 | Measurement tools including: distance, depth, area, and circumference | #
3.23 | Volume Flow Measurements | #
3.24 | On-board workstation-class data management with thumbnail previews and storage of images, loops, and reports | #
3.25 | Retrospective and prospective clip capture to internal drive or removable media | #
3.26 | Integrated DVD/CD burning capability for storage of DICOM 3 images or export in JPEG and avi for PC compatibility | #
3.27 | DICOM 3.0 Print and Store capability to internal drive or DVD/CD | #
3.28 | Internal HDD at least 1 TB | 2
3.29 | One Full Off-line Work station with Full Reporting Systems | #
3.3 | Chair, ergonomic, height adjustable with lockable castors for the Health professional. | 2
3.31 | Gel warmer | 2
3.32 | Foot rest | 1
3.33 | Cable tray /hooks ..cable support system | 1
3.34 | Easy removable air filters | 2
3.35 | USB PORT additional | 2
4 | TRANSDUCERS
4.1 | The following transducers below must be included as standard. Details are stated:
4.1.1 | Adult Sector transducer (1 MHz Hz to 5 Mhz) using innovating single crystal technology. | #
4.1.2 | Paediatric Sector transducer (4 Mhz to 12 MHz) using innovating single crystal | #
4.1.3 | Neonatal Sector transducer (2 Mhz to 8 MHz) using innovating single crystal | #
4.1.4 | Matrix transducer for transthoracic 4D / Live 3D Echo | #
4.1.5 | Matrix Trans-esophageal transducer (TEE) using matrix array technology for 4D / Live 3D trans-esophageal Imaging | #
4.1.6 | Matrix Trans-esophageal transducer (TEE) using Single crystal technology for 2D imaging | #
4.1.7 | Matrix Paediatric Trans-esophageal transducer (TEE) using Single crystal technology for 2D imaging | #
4.1.8 | Linear vascular Carotid transducer Probe frequency between 4-6 MHz lower range and 12 to 14 MHz upper range | #
4.1.9 | Non-Imaging Pencil Probe with 1.9 MHz and above | #
4.2 | The following transducers must be included as optional. The end user must be able to make the final selection of transducers. Clearly State all available transducers with prices per transducers (swapping of products/transducers and software to be allowed with equitable costs). Details are stated.
4.2.1 | Convex Probe (Curvilinear) with tissue harmonics. The MHz range are stated (within 1-6 MHz) | #
4.2.2 | Microconvex for neonatal / paediatrics with tissue harmonics.The MHz range are stated (4-8) | #
4.2.3 | Microconvex Transvaginal / transrectal with tissue harmonics. The MHz range are stated (within 3-8 MHz) | #
4.2.4 | Linear probe (vascular) with tissue harmonics. The MHz range are stated (4-12) | #
4.2.5 | Linear probe (MSK) with tissue harmonics.The lower MHz range must be 5-6 MHz and the upper range must be between 14 and 21 MHz. | #
4.3 | New active matrix array technnology on transthoracic and Transesophageal transducers | 1
4.4 | The selected frequency of the imaging transducer must be displayed on the screen. | #
5 | GENERAL
5.1 | Small Black White photo printer integrated into the system | #
5.2 | Measurement and Report software packages | #
5.3 | Minumum of 3-Lead ECG Cable with connecting port | #
5.4 | Integrated spyrometric function(display respiratory motion) | #
7 | Optional
7.1 | Ultrasound Quality Control Phantom (swapping of phantoms, products and software to be allowed with equitable costs). Details are stated. | 2
7.2 | side storage bin
7.3 | A second Full Off-line Work station with Full Reporting System as optional item | 2
7.4 | External Hardrive of 1 TB as optional item | 2
8 | POWER SUPPLY
8.1 | The unit must be supplied with an on-line uninterrupted power supply (U.P.S.) to provide minimum ten 10 to 20 minutes mains backup power.
8.2 | The UPS should not adversely affect the image quality.
8.3 | This unit must be included in the bid price and be covered by the guarantee and maintenance plan.