## Sheet1
ULTRASOUND MACHINE, HIGH END WITH TEE
ITEM | TECHNICAL SPECIFICATIONS | Weight
1 | Ultrasound high end with TEE
1.1 | Unit should offer at least a 17 to 19 inch LCD high resolution colour monitor. State the size and technology. | 2
1.2 | Unit must be latest technology. two year of which the unit was manufactured. | 2
1.3 | The Unit must be a Fully Digital high end colour doppler ultrasound unit with excellent colour and power sensitivity while maintaining excellent detail resolution. | 2
1.4 | Unit must be housed on a lightweight mobile cart with lockable wheels. | 2
1.7 | The Width of the system in cm 100 to 156cm | 2
1.8 | The Depth of the system must be 100 to 156 of the system in cm | 2
1.9 | User configurable hard keys | 2
1.1 | User interface must be height adjustable | #
1.11 | the maximum number of channels 512 | 2
1.13 | the system dynamic range must be 45 to 90db | #
1.15 | Unit should offer hard drive with image storage space of 300GB to 500GB . | 2
1.16 | CD/DVD writer with the required software for active use, must be standard and integrated in the system | #
1.17 | Recorded data on CD/DVD must be readable on any standard laptop/computer with additional software, if necessary from the supplier | #
1.18 | The system must be DICOM compliant | #
1.19 | Advanced DICOM with full worklist, transmission and print must be offered and be included in main system | #
1.2 | Export of images from the system must be possible. (including transfer of images to PACS). State the methods. | #
1.21 | The system must have a touch command screen for ease of operations | #
1.22 | Digital black and white thermal printer must be offered as part of the system. | #
2 | Imaging Modes
2.1 | Unit should be capable of offering a maximum frame rate of not less than 500 frames per second. | #
2.2 | The Maximum scanning depth of the abdominal convex probe/transducer must be 40cm and above . State the maximum scanning depth of the range of probes provided | #
2.3 | Independent gain control for Pulsed Wave | #
2.4 | The PW sampling range 0.2 to 20mm (gate sizes) | 2
2.5 | State the Pulse Repetition Frequency 10000 to 19000Hz. | 2
| Streerable continuous wave repetition frequency | 2
2.7 | State the Continuous wave repetition frequency 1.56 to 34.7Khz . | 2
2.8 | detectable velocity for CW Doppler to be 30 to 650cm per second | 2
2.9 | 2D image on linear transducers must be electronically steerable. The are up to 7 steering angles available on linear transdusers | 1
2.1 | State the degree of Independent angle steering. | 1
2.11 | Post processing must be available on frozen images | #
2.12 | Independent M-Mode gain adjustment must be available | #
2.13 | Indicate if the unit are able to support Elastography Software to display the elasticity of tissues in real time | 1
2.14 | Indicate if the unit are able to support a software that depicts small calcifications. | 1
2.15 | Indicate if the unit are able to support Panoramic Software. If so, state the maximum length on offer. | 1
2.16 | Indicate if the unit are able to offer trapezoid sofware capabilities on all sector and linear probes. | 2
2.17 | The Grey scale level must be 256 as a minimum. | 2
2.18 | maximum amount of images that can be stored on the hard drive must be 200 000 to 300 000 images. | 2
2.19 | State the maximum storage capability in GB that is available on your hard drive. | 2
2.2 | Conventional Tissue Harmonics must be offered on all probes. | #
2.21 | Differential Tissue Harmonics to improve spatial resolution and tissue visualization must be offered. | 2
2.22 | Indicate if the unit are able to offer pulse substraction technique. | 1
2.23 | 2D Loop up to 400 frames should be recorded. State the total frames if ˃ than 400. | 1
2.24 | Indicate the maximum length of the captured loop in minutes. | 2
2.25 | Real time triplex mode must be standard. | #
3 | Doppler and Colour Doppler
3.1 | State the different Doppler display modes i.e.. side by side or top to bottom | 2
3.2 | All probes must offer Pulsed Wave Doppler. | 2
3.3 | State the maximum detectable velocity for Pulsed Wave. The maximum detectable velocity for pulse wave 350 - 500 cm/per sec. | 2
3.4 | Image memory to operate in both dual and quad screen format. | 2
3.6 | State if simultaneous display of the colour image and 2D image, side by side, in real-time is possible. | 2
3.7 | Indicate the range of independent colour gain adjustment in 1 dB steps. | 2
3.8 | The colour frame rate should be 400 or above. State maximum colour frame rate. | 2
3.9 | Colour flow must be fully operational through the total depth of the image. | #
3.1 | On screen measurements must be possible from real time examination. | #
3.11 | Unit must offer user programmable annotation for quick labelling. | #
3.12 | Storage and review of patent data must be available to compare exams. | #
3.13 | Indicate the maximum number of processing selections available, for example, wall filters, colour averaging. List the processing selections available.Indicate if the user are able to store them in presets. | 2
3.14 | Indicate if Tissue Doppler Imaging (TDI) for cardiac application are standard on B-Mode and M-Mode in the unit. | 2
3.15 | State type of remote control device that links with system during difficult exams. (Please quote optionally) | 1
3.16 | Controlling the system remotely with a remote control | 1
3.17 | Controlling the system remotely with voice commands | 1
3.18 | Controlling the system remotely by any other method. Explain method if applicable. | 1
3.18.1 | Indicate if Contrast harmonics software are be available. Quote optionally | 1
3.18.2 | Indicate if Advanced Dynamic Flow are offered on all transducers for vascular recognition imaging | 2
3.19 | Indicate if Directional Colour Angio are available to display the existance of blood flow and the direction of flow movement with high sensitivity | 2
4 | Transducers/ Probes
4.1 | The selected frequency of the imaging transducer must be displayed on the screen. | #
4.2 | Five transducers must be included and the end user must be able to select the final transducers. Clearly State all available transducers with prices per transducers (swapping of products/transducers and software to be allowed with equitable costs). Details are stated. | 2
4.3 | Additional probes and software must be available as optional items. Optional items must be listed and priced separately. | 2
4.4 | Types and frequency range of Transducers
4.4.1 | Convex Probe (Curvilinear) with tissue harmonics. The MHz range are stated (within 1-6 MHz) | #
4.4.2 | Microconvex for neonatal / paediatrics with tissue harmonics.The MHz range are stated (4-8) | #
4.4.3 | Microconvex Transvaginal / transrectal with tissue harmonics. The MHz range are stated (within 3-8 MHz) | #
4.4.4 | Linear probe (vascular) with tissue harmonics. The MHz range are stated (4-12) | #
4.4.5 | Linear probe (MSK) with tissue harmonics.The lower MHz range must be 5-6 MHz and the upper range must be between 14 and 21 MHz. | #
4.4.6 | Adult Cardiac Probe with tissue harmonics. State the MHz range (1-4MHz). | #
4.4.7 | Paediatric Cardiac Probe. State the MHz range (2-10 MHz). | #
4.4.8 | Neonatal Cardiac Probe. State the MHz range (2-10 MHz). | #
4.4.9 | Mono-plane Endorectal/ (transrectal ) transducer with tissue harmonics . State the MHz range. | 2
4.4.10 | Biplane Endorectal (transrectal) transducer, linear and linear , with tissue harmonics . State the MHz range. | 2
4.4.11 | Biplane Endorectal (transrectal)transducer, linear and convex , with tissue harmonics . State the MHz range. | 2
4.4.12 | Transesophageal echocardiography (TEE)with tissue harmonics. State the MHz range. | #
4.4.13 | Paediatric Transesophageal echocardiography (TEE)with tissue harmonics. State the MHz range. | #
4.4.14 | State your advanced cardiac applications as an option. | 2
4.4.15 | State your Speckle imaging application as an option | 1
4.4.16 | State your 2D, 3D and 4D MPR applications as options | 1
4.4.17 | Tissue harmonics should be available on all offered transducers | #
4.4.18 | Indicate the footprint size of the Pediatric sector transducer | 1
4.4.19 | The selected frequency of the imaging transducer must be displayed on the screen. | #
4.4.20 | All transducers must be digitally controlled so that continuous beam forming is possible | #
4.4.21 | System must have 4 or more active imaging transducer connectors with electronic switching between probes. | 2
5 | Measurements
5.1 | 16 callipers (8channel) for measuring distance/volume/angle and area must be offered | 2
5.2 | 16 callipers (8 channels) hip joint should be offered | 2
5.3 | State the storage facility offered for patient fetal growth measurement data files. The information must be able to be retrieved for follow-up comparison. 250GB must be offered. | 2
5.4 | Basic Measurements: -B Mode: Distance , area, angle, volume, stenosis ratio; 4D mode basic measurements: Distance, Volume; M-Mode, basic measurement: Slope, Distance, Time, Heart Rate; PW/CW Mode Measurements: acceleration, velocity, time,heart rate, PI,RI, S/D,Flow Volume, Doppler Trace | #
5.5 | Unit must offer 50th percentile growth graphs from the Doppler measurements, RI/PI. The information must be able to be retrieved for follow-up comparison | 2
5.6 | 50th percentile growth graphs must be offered on all measurements for example BPD, FL, AC, Fetal weight, HC. The information must be able to be retrieved for follow-up comparison | #
5.7 | Display of graphs must be available to compare measurements of twins or triplets in single and quad display format | #
5.8 | System must offer user programmable annotation for quick labelling | #
5.9 | Applications measurements should automatically be labelled/displayed on the touch command screen for all applications, (2D mode and Doppler). Comprehensive reports must be generated automatically. | 2
5.1 | Chair, ergonomic, height adjustable with lockable castors for the Health professional. | 2
5.11 | Foot rest as a standard | 1
5.12 | Cable tray /hooks ..cable support system | 1
5.13 | easy removable air filters | 2
5.14 | Gel warmer. | 2
5.15 | USB PORT additional | 2
6 | Recording devices
6.1 | Digital colour thermal printer must be offered as a standard. Printer mounted on the machine | 1
7 | POWER SUPPLY
7.1 | The unit must be supplied with an on-line uninterrupted power supply (U.P.S.) to provide minimum ten 10 to 20 minutes mains backup power. | #
7.2 | The UPS should not adversely affect the image quality.
7.3 | This unit must be included in the bid price and be covered by the guarantee and maintenance plan.
8 | OPTIONAL ITEMS
1 | fold down monitor for easy transportation | 1
| Measurement and Report software packages | 2
| Details of tools/software provided as a standard and those as optional, are stated | 1
| side storage bin | 1
| Ultrasound Quality Control Phantom (swapping of phantoms, products and software to be allowed with equitable costs). Details are stated. | 2