## Sheet1
Breathing Simulator for Ventilatory Management Training
1 | DESCRIPTION
1.1 | Breathing simulator for basic to advanced ventilatory management training across all modes and types of ventilation
2 | TECHNICAL SPECIFICATION
2.1 | Complete Breathing Simulator for Ventilatory Management Training with the following:
2.1.1 | Breathing simulator operating system
2.1.2 | Lung adaptor for full-body, high fidelity manikin(s)
2.1.3 | Lung circuit adaptor
2.1.4 | Lung breathing circuit
2.1.5 | O-rings and seals
2.1.6 | Gas scrubbing system (optional)
2.2 | A breathing simulator that can connect to any ventilator and elicit the same breathing dynamics as a real patient across a broad spectrum of normal and pathological respiratory conditions and manifestations
2.3 | Presents age-appropriate respiratory dynamics from birth to old age
2.4 | Accurate physiological modelling of respiratory patterns
2.5 | Breathe-to-breathe changes of lung dynamics using purpose-designed software as a stand-alone system or integrated into manikin-based simulation
2.6 | Lung mechanics features:
2.6.1 | Must simulate the following lung mechanics
2.6.2 | Adjustable compliance ranges from 0.5-250 mL/cmH2O
2.6.3 | Adjustable resistance ranges from 8-150 cmH2o/L/s
2.6.4 | Generate a tidal volume for spontaneous breathing ventilation support from 2 to 2500 ml
2.7 | Ventilator parameter interaction:
2.7.1 | Programmable spontaneous breathing rate from 0-100bpm
2.7.2 | Programmable spontaneous breathing effort from 0-100cmH2O
2.7.3 | Programmable Inspiratory : Expiratory (I:E) ratio
2.7.4 | Demonstrate treatment effects of Positive End Expiratory Pressure (PEEP) across clinically relevant settings including PEEP values > 20cmH20
2.8 | Case development features for ventilatory management:
2.8.1 | Support non-invasive ventilation strategies such as Bag-valve Mask Ventilation, Continuous Positive Airway Pressure (CPAP) and Bilevel Positive Airway Pressure (BiPAP)(High flow nasal)
2.8.2 | Support all modes of ventilation including: Pressure/Volume Control, Pressure Support, Airway Pressure Release Ventilation (APRV), Adaptive Support Ventilation (ASV), Proportional Assist Ventilation (PAV), and High Frequency Oscillatory Ventilation (HFOV)
2.8.3 | Pre-programmed settings for respiratory conditions such as asthma, Acute Respiratory Distress Syndrome (ARDS), Chronic Obstructive Pulmonary Disease (COPD) and Interstitial Lung Disease (ILD) with variable levels of severity
2.8.4 | Programmable user-defined respiratory conditions such as pneumothorax, bronchospasm, pneumonia, cystic fibrosis (as examples)
2.8.5 | Program and simulate advanced ventilation presentations / management strategies such as patient-ventilator dyssynchrony, weaning and waveform analysis
2.9 | System Integration / compatibility requirements:
2.9.1 | Can be upgraded as a breathing simulator for premature babies
2.9.2 | Compatible with high fidelity, full body adult and infant manikins
2.9.3 | Lung software must integrate with the manikin operating system (clarify the compatibility with the end user's system)
2.1 | Must come with a product user guide
2.11 | Must come in a protective case
Breathing Simulator for Ventilatory Management Training - Accessories and Consumables
1 | Breathing simulator operating system
2 | Mechanism for generating tidal volume
3 | Resistance mechanism
4 | Compliance mechanism
5 | Lung adaptor for full-body, high fidelity manikin(s)
6 | Adaptor module for simulating lung dynamics of premature babies
7 | Lung circuit adaptor
8 | Lung breathing circuit
9 | O-rings and seals
10 | Gas scrubbing system (optional)