## Sheet1
High Fidelity Male Manikin
1 | TECHNICAL SPECIFICATION
1.1 | The manikin must conform to the following GENERAL specifications:
1.1.1 | Must be adult life-like male size manikin
1.1.2 | Must be wireless and tetherless and must be functional for prehospital transportation scenarios, which includes Bluetooth capabilities
1.1.3 | Must be able to work off main power supply and internal power supply (batteries)
1.1.4 | All required software and hardware must be supplied for full functionality
1.1.5 | Must have case-based ultrasound capabilities for virtual or real-life monitoring (fast, e-fast, and rush protocols)
1.1.6 | Must integrate or be compatible with capture and storage of audio, visual and technical data capture for debriefing, education and research
1.1.7 | Must be available in various skin colours (light, medium, and dark tones) as per end user requirements
1.1.8 | OPTIONAL: The manikin must be supplied with all required interchangeable parts to change the manikin to a female patient if so required.
1.2 | The manikin’s skin must have lifelike features with articulation on the following areas:
1.2.1 | Neck
1.2.2 | Jaw
1.2.3 | Shoulders
1.2.4 | Elbow
1.2.5 | Hips
1.2.6 | Knees
1.2.7 | Ankles
1.3 | The manikin must also be fully articulating for placement of the following common patient positions:
1.3.1 | Supine
1.3.2 | Lateral
1.3.3 | Fowler’s position
1.4 | The manikin must conform to the following AIRWAY capabilities and specifications:
1.4.1 | Must have anatomically accurate oral cavity and airways
1.4.2 | Must be able to take advanced airways and must detect advanced airway placement on programming:
1.4.2.1 | Endotracheal intubation (with right mainstem intubation detection)
1.4.2.2 | Supraglottic airway (LMA)
1.4.2.3 | Oropharyngeal airway
1.4.2.4 | Nasopharyngeal airway
1.4.3 | Must have electronic adjustments to simulate difficult airways
1.4.3.1 | Must include the tongue edema, laryngospasm
1.4.4 | Must include “can’t intubate can’t ventilate” scenarios
1.4.5 | Must be able to perform surgical airway techniques on the manikin
1.4.6 | Must be able to perform suctioning techniques on the above airways.
1.4.7 | Must be able to perform basic airway opening techniques
1.4.7.1 | jaw thrust
1.4.7.2 | head-tilt-chin-lift
1.5 | The manikin must conform to the following BREATHING capabilities and specifications:
1.5.1 | Normal and Abnormal sounds must be audible when using spontaneous ventilations or during assisted ventilations
1.5.2 | Must have lifelike, spontaneous breathing which can be adjusted using simulation programming to be either unilateral or bilateral chest rise
1.5.3 | Must have multiple anterior and posterior auscultation sounds bilaterally
1.5.4 | Must include a library of normal and abnormal auscultation sounds
1.5.5 | Must be able to support the use of a real mechanical ventilator, with or without the requirements of an adapter with use of an external specialised lung
1.5.6 | The manikin must have the following realistic breathing adoptions:
1.5.6.1 | Variable lung compliance
1.5.7 | Manikin must make provision for chest interventions:
1.5.7.1 | Hemo/pneumothorax site
1.5.7.2 | Chest tube insertion
1.5.7.3 | For needle thoracentesis:
1.5.7.3.1 | Must support needle insertion
1.5.7.3.2 | Must provide audible hiss on correct insertion
1.6 | The manikin must conform to the following CIRCULATION capabilities and specifications:
1.6.1 | Must have palpable pulses, bilaterally on the following sites:
1.6.1.1 | Carotid
1.6.1.2 | Radial
1.6.1.3 | Femoral
1.6.1.4 | Pedal
1.6.2 | All pulses must be recorded on the manikin management program
1.6.3 | The following skin changes must be performable:
1.6.3.1 | Cyanosis
1.6.4 | Must be able to use real or virtual vital signs monitoring that integrates with manikin function
1.6.5 | Must be able to support real-time intravenous fluid/infusion therapy.
1.6.6 | Must be able to realistically perform phlebotomy
1.6.7 | Must be capable of performing intraosseous insertion and fluid administration
1.6.8 | Must be capable of performing intramuscular injections
1.7 | The manikin must conform to the following NEUROLOGICAL capabilities and specifications:
1.7.1 | Adjustable blinking eyes, Pupil reaction
1.7.2 | Pupils must realistically adjust to light, but must be programmable
1.7.3 | The simulation controller must have pre-recorded responses
1.7.4 | The neck of the manikin must be able to simulate the following:
1.7.4.1 | Rotation
1.7.4.2 | Flexion
1.7.4.3 | Extension
1.7.4.4 | Reduced movement during neck injuries
1.7.5 | Must have appropriate movement representing seizure activity
1.7.6 | Must be able to simulate sweating and tears.
1.8 | The manikin must conform to the following GENITOURINARY capabilities and specifications:
1.8.1 | Must be able to insert urinary catheter
1.8.2 | Must be supplied with an internal reservoir for urine.
1.9 | The manikin must conform to the following GASTROINTESTINAL capabilities and specifications:
1.9.1 | Bowel auscultation field (minimum 2 fields) with a library of various sounds
1.9.2 | Visible gastric inflation
1.1 | The manikin must conform to the following TRAUMA capabilities and specifications:
1.10.1 | Must have dedicated ports for bleeding simulation and detachable trauma limbs
1.10.2 | Must include traumatic module set
1.10.3 | Must have an internal reservoir to store blood
1.11 | The manikin must conform to the following CARDIAC capabilities and specifications:
1.11.1 | Auscultation of heart sounds must be realistic on Aorta, Pulmonary, Tricuspid, and mitral valves of the heart
1.11.2 | Must be able to realistically place 4 lead ECG wires
1.11.3 | Must be able to program and simulate cardiac injuries which include:
1.11.3.1 | Ischemia
1.11.3.2 | Infarction
1.11.4 | Must include a 12-lead library
1.11.5 | Must support the following cardiac interventions:
1.11.5.1 | Defibrillation
1.11.5.2 | Cardioversion
1.11.5.3 | Pacing
1.11.6 | Must allow defibrillation with paddle or pad placement as either:
1.11.6.1 | Anterior / Lateral
1.11.7 | Manikin and software must have a feedback mechanism on compressions (rate, depth interruption time), and respirations (excessive ventilation, rate)
1.12 | Must come with a product user guide
1.13 | Must come in a protective case
High Fidelity Male Manikin - Accessories and Consumables
1 | artificial blood
2 | lubricant
3 | artificial urine
4 | talcum powder