## Sheet1
High Tech High Fidelity Maternal Manikin
1 | TECHNICAL SPECIFICATION
1.1 | The manikin must conform to the following GENERAL specifications:
1.1.1 | Must be adult life-like female size manikin
1.1.2 | Must be wireless and tetherless and must be functional for prehospital transportation scenarios, which includes Bluetooth capabilities
1.1.3 | Must be able to work off main power supply and internal power supply (batteries)
1.1.4 | All required software and hardware must be supplied for full functionality
1.1.5 | Must integrate or be compatible with capture and storage of audio, visual and technical data capture for debriefing, education and research
1.1.6 | Must be available in various skin colours (light, medium, and dark tones) as per end user requirements
1.1.7 | Must be fully articulating with realistic skin and facial features
1.1.8 | The manikin must be able to convert into a non-pregnant patient for general gynecological training
1.2 | The manikin’s skin must have lifelike features with articulation on the following areas:
1.2.1 | Neck
1.2.2 | Jaw
1.2.3 | Shoulders
1.2.4 | Elbow
1.2.5 | Hips
1.2.6 | Knees
1.2.7 | Ankles
1.3 | The manikin must also be fully articulating for placement of the following common patient positions:
1.3.1 | Supine
1.3.2 | Prone
1.3.3 | Fowler’s position
1.4 | The manikin must conform to the following AIRWAY capabilities and specifications:
1.4.1 | Must have anatomically accurate oral cavity and airways
1.4.2 | Normal and Abnormal sounds must be audible when using spontaneous ventilations or during assisted ventilations
1.4.3 | Must be able to take advanced airways and must automatically detect advanced airway placement:
1.4.3.1 | Endotracheal intubation (with right mainstem intubation detection)
1.4.3.2 | Supraglottic airway (LMA)
1.4.3.3 | Oropharyngeal airway
1.4.3.4 | Nasopharyngeal airway
1.4.4 | Must have electronic adjustments to simulate difficult airways (e.g., Tongue edema, Laryngospasm, Pharyngeal swelling)
1.4.5 | Must include “can’t intubate can’t ventilate” scenarios
1.4.6 | Must be able to perform suctioning techniques on the above airways.
1.4.7 | Must support the use of a Bag Valve Mask resuscitator and mechanical ventilations
1.5 | The manikin must conform to the following BREATHING capabilities and specifications:
1.5.1 | Must have lifelike, spontaneous breathing which can be adjusted using simulation programming to be either unilateral or bilateral chest rise
1.5.2 | Must have multiple anterior and posterior auscultation sounds bilaterally
1.5.3 | Must include a library of normal and abnormal auscultation sounds
1.5.4 | Must be able to support the use of a real mechanical ventilator, with or without the requirements of an adapter with use of an external specialised lung
1.5.5 | The manikin must have the following realistic breathing adoptions:
1.5.5.1 | Real CO2 exhalation module with capnography waveforms must be supplied
1.6 | The manikin must conform to the following CIRCULATION capabilities and specifications:
1.6.1 | Must have palpable pulses, bilaterally on the following sites:
1.6.1.1 | Carotid
1.6.1.2 | Brachial
1.6.1.3 | Radial
1.6.2 | All pulses must be recorded on the manikin management program
1.6.3 | Must be able to use real or virtual vital signs monitoring that integrates with manikin function
1.6.4 | Must be able to support real-time intravenous fluid/infusion therapy.
1.6.5 | Must be able to realistically perform phlebotomy
1.6.6 | Must automatically recognize medication administration
1.6.7 | Must be capable of performing intramuscular injections and sub cutaneous injections
1.7 | The manikin must conform to the following NEUROLOGICAL capabilities and specifications:
1.7.1 | Adjustable blinking eyes, Pupil reaction, and eye movement
1.7.2 | Pupils must realistically adjust to light, but must be programmable
1.7.3 | Must have realistic eye movement with object tracking
1.7.4 | Eyes must be able to simulate the following:
1.7.4.1 | Strabismus
1.7.4.2 | Nystagmus
1.7.4.3 | Blepharospasm
1.7.4.4 | Ptosis
1.7.4.5 | Mydriasis
1.7.4.6 | Anisocoria
1.7.5 | The simulation controller must have pre-recorded responses
1.7.6 | The neck of the manikin must be able to simulate the following:
1.7.6.1 | Rotation
1.7.6.2 | Flexion
1.7.6.3 | Extension
1.7.7 | Must have appropriate movement representing seizure activity
1.8 | The manikin must conform to the following GENITOURINARY capabilities and specifications:
1.8.1 | Must be able to insert urinary catheter
1.8.2 | Urinary flow must be programmable
1.8.3 | Must be supplied with an internal reservoir for urine.
1.9 | The manikin must conform to the following GASTROINTESTINAL capabilities and specifications:
1.9.1 | Bowel auscultation field (minimum 4 fields) with a library of various sounds
1.9.2 | Rectum must support the use of suppository placements
1.1 | The manikin must conform to the following CARDIAC capabilities and specifications:
1.10.1 | Auscultation of heart sounds must be realistic on Aorta, Pulmonary, Tricuspid, and mitral valves of the heart
1.10.2 | Must be able to realistically place 4 lead ECG wires
1.10.3 | Must be able to program and simulate cardiac injuries which include:
1.10.3.1 | Ischemia
1.10.3.2 | Infarction
1.10.3.3 | Necrosis
1.10.4 | Must include a 12-lead library or must be able to generate a 12 lead ECG through software (must be supplied)
1.10.5 | Must support the following cardiac interventions:
1.10.5.1 | Defibrillation and double sequential defibrillation
1.10.5.2 | Cardioversion
1.10.5.3 | Pacing
1.10.6 | Must allow defibrillation with paddle or pad placement as either:
1.10.6.1 | Anterior / Lateral
1.10.7 | Manikin and software must have a feedback mechanism on compressions (rate, depth interruption time), and respirations (excessive ventilation, rate)
1.11 | The manikin must conform to the following OBSTETRIC capabilities and specifications:
1.11.1 | The mother and baby must be able to physiologically link
1.11.2 | Must support fetal monitoring using a real device
1.11.3 | Must be supplied with and support a fetal delivery system for realistic simulation of fetal descent and movements
1.11.4 | Must realistically simulate the following obstetrical emergencies:
1.11.4.1 | Breech delivery
1.11.4.2 | Shoulder dystocia
1.11.4.3 | C-section deliveries
1.11.5 | The birth canal must be self-lubricating and computer controlled
1.11.6 | Must be supplied with realistic placenta and include various detachable fragments, and umbilical cord
1.11.7 | Must support the use of Epidural needle placement with detection of insertion. Landmarks must be realistic
1.11.8 | Must simulate intrapartum bleeding
1.11.9 | Must indicate, through sensors, traction and torsion applied on fetus
1.11.10 | Must be able to perform:
1.11.10.1 | McRoberts maneuver
1.11.10.2 | Suprapubic pressure
1.11.10.3 | Zananelli
1.11.10.4 | Wood’s screw
1.11.10.5 | Hands and knees position
1.11.11 | The fetus must be supplied and must be a realistic size for a full-term neonate
1.11.12 | Must include replaceable episiotomy inserts
1.11.13 | Episiotomy must simulate bleeding and tearing
1.11.14 | The uterus must be programmable for firmness, and must support hemorrhaging and balloon tamponade
1.11.15 | Hemorrhaging must be controlled through computer
1.11.16 | Must be able to simulate and realistically manage post-partum hemorrhage with automatic deterioration of vital signs from blood loss
1.12 | The neonate for this manikin must be supplied and conform to the following specifications:
1.12.1 | Must be the size of a full-term neonate
1.12.2 | Must be able to simulate the following:
1.12.2.1 | Heart rate
1.12.2.2 | Respiratory rate
1.12.2.3 | Crying
1.12.2.4 | Cyanosis
1.12.2.5 | Movement
1.13 | The simulator must be supplied with all hardware and computer software compatible with the manikin and must include:
1.13.1 | A minimum of 9 labor and delivery simulations compatible with this manikin must be supplied
1.13.2 | Must be supplied with an interchangeable deep vein thrombosis leg that matches the colour of the patient.
1.14 | Must come with a product user guide
1.15 | Must come in a protective case
High Tech High Fidelity Maternal Manikin - Accessories and Consumables
1 | artificial blood
2 | lubricant
3 | artificial urine
4 | talcum powder