## Sheet1
Diagnostic Audiometer (headphones and inserts), including: air and bone conduction, speech audiometry, masking
Description
Audiometer used to test and diagnose hearing thresholds.
Technical Specification
Clinical, two-channel audiometer that must test in increments of no more than 5db.
Capable of performing masked and unmasked pure tone and warble tone air and bone conduction and speech testing
Masking sounds must include narrowband, white and speech noise
Frequency range for air conduction must be within a range of 125Hz to 8000Hz, with intensity range of -10dB to 120dB.
The intensity range for bone conduction to span between -10dB to 80dB for frequencies of 250-4000Hz.
Must provide for contralateral white noise, speech noise and narrow band noise for air / bone and free field audiometry must be provided on the unit offered
The audiometer offered must allow for user selectable attenuation over the entire range
Must have the option to add following special tests:
SISI
ABLB
stenger
bekesy
weber
tone-in-noise
speech-in-noise
Must come with:
earphones
paired insert phones
bone conductor
response button
talk-forward
built-in or gooseneck microphone or operator's headset with microphone
Must provide auditory as well as visual patient response monitoring and intensity controls in order to further enhance user control of its operation.
Pre-recorded speech wordlists in English (with the option of other South African languages) in male and female voice can be built-in, failing which the audiometer must have a built-in CD player or must have a USB connection and come with a compatible external CD Player
Accessories:
earphones
paired insert phones
bone conductor
response button
talk-forward
built-in or gooseneck microphone or operator's headset with microphone
Optional accessories:
field speakers
SISI
ABLB
stenger
bekesy
weber
tone-in-noise
speech-in-noise
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list