## Sheet1
Portable Diagnostic Audiometer and Tympanometer with extended high frequencies (tele-audiometry)
Description
Portable audiometer tympanometer used for testing and diagnosing hearing thresholds including suprathresholds level testing that allows for tele-audiometry capabilities.
Technical Specification
Two-channel audiometer that must test a range of frequencies from 125-16 000Hz, with intensities of -10 to 110dB, in increments of no less than 5db
Must be capable of performing masked pure-tone and warble-tone air and bone conduction and speech testing
Must come with a with:
PC/laptop-integrated or controlled system
protective carry case
earphones
paired insert earphones
bone conductor
response button
talk-forward
built-in or gooseneck microphone
Earphones must have individual audiometers with active noise monitoring technology to allow diagnostic testing without a soundproof booth
Must have a built-in pre-recorded speech material in South African accents and languages or USB ports to connect to an external CD player (to be included if no built-in pre-recorded speech material) for recorded tests
Must allow for tele-audiometry
Must allow for automated and manual testing
Must allow for printing capabilities
Accessories
protective carry case
earphones
paired insert earphones
bone conductor
response button
built-in or gooseneck microphone
Technical Specifications of screening tympanometer (226Hz tone):
Tympanometry measurement capability of 226Hz.
Must test tympanometry across range of +200 -to -400daPa with accuracy of ±10daPa
Ipsi-lateral testing range from 500,1000,2000 and 4000Hz
Must include probe/ear tips
Must be supplied with protective carry case
Must have A/C power supply
Must have printing capabilities and supplied with two rolls of paper
Must have a memory for at least 12 tests, with probe assembly
Must be supplied with a box of ear tips of at least six different sizes, with at least three tips of each size
Accessories:
probe/ear tips of at least six different sizes, with at least three tips of each size
protective carry case
two rolls of paper
memory for at least 12 tests
with probe assembly
Box of ear tips
Optional accessory
Printer
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list