## Sheet1
Screening Oto-Acoustic Emission and Automated Brainstem Response combo
Description
Portable and rechargeable oto-acoustic emission and Automated brainstem response (OAE and ABR) system allowing for the screening of outer hair cell functioning and auditory brainstem response.
Technical Specification
Technical specifications for Screening OAE:
Must be capable of measuring DPOAEs with frequencies: 2,3,4,5 kHz and intensity: 65/55 dB
With an option of measuring TEOAE's with frequencies: 1.5-4 kHz and intensity: 83 dB SPL
Must have predefined and user defined protocols
Should be able to connectable to a PC or printer
Must be supplied with:
carry case
electrode probes or pads for the ABR
various ear tips sizes for the OAE including neonatal sizes
USB cable to connect to PC
probe case
Must provide electrodes and electrodes pads or insert earphones
Must include tips of various sizes including neonates
Accessories:
ear tips
electrode pads
Technical specifications for ABR:
Must have the following parameters:
stimulus types: Click 90-100 µsec
ABR waveform capture, display and archives tests
Stimulus Rate: 40 to 60 per second
Stimulus Intensity: 35dB
input frequency Range: 30Hz to 3,000Hz
waveform latency delay: ±0.4 msec, Gain: 40-45 dB
Should contain integrated electrodes pads or insert earphones.
Must be able to be connected to a printer for printing purposes, with a rechargeable printer
Provide inserts of various sizes including neonates.
Laptop requirements are a minimum of 5GB free disk space, minimum 1GB of RAM, latest Intel version
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list