## Sheet1
Soundproof Booth with Extraction Fan (small)
Description
Soundproof booth with ventilation and extractor fan (ventilation unit for TB protection with minimal noise output).
Technical Specification
Screening soundproof booth with ventilation and extractor fan (with minimal noise output).
Should contain observation safety glass window and work platform attachment.
The booth should have sound absorption in accordance with ISO 9001, SANS or equivalent.
Inside ambient noise levels compliant with the ANSI S.1-1999 or equivalent.
Fire resistance in compliance with EN13823 and SANS 10177.
Minimum external dimensions of length: 975mm, width: 735mm, height 1917mm.
Minimal internal dimensions of length: 850mm, width: 610mm, height: 1720mm.
At least 6 input and output jacks (for air conduction left and right, bone conduction, talk forward, and talk back)
Should contain observation safety glass window which is appropriately acoustically sealed and work platform attachment which is appropriately sized to accommodate audiometer, accessories and a writing surface.
The booth should have sound absorption in accordance with ISO 9001 or equivalent
The proposed audiology booth is to comply with SANS 10182: 2006 as amended– the measurements and assessment of acoustic environments for audiometric test.
The booth must be mounted on anti-vibration mounts to avoid any structural transfer of sound from the main structure.
Inside ambient noise levels compliant with the ANSI S.1-1999 or equivalent.
Fire resistance in compliance with EN13823 and SANS 10177.
Should have adequate silent ventilation system that does not generate sound in the booth or conduct sound through the envelope of the booth resulting in an ambient sound pressure level above the prescribed limit per octave band
Outward opening door that can be opened from inside as well as outside of the booth, is not lockable from the inside, is at least 900 mm wide, and must appropriately acoustically seal so as not to compromise the required ambient sound pressure level inside the booth
Comes standard with wheelchair access ramp that does not impede door closure
Input jacks for connecting equipment on the outside of the booth
Output jacks for connecting transducers inside the booth
Adequate lighting supplied within the booth and clinician test area when appropriate.
Electrical fittings and wiring to have standard certified safety measures, including trip switches.
All penetrations through the envelope for electrical services must be sealed and must not compromise the ambient sound pressure level inside the booth.
On/off switches for lighting and ventilation to be installed in easy to reach location for audiologist.
Two standard 3 pin electrical outlet to be installed at suitable location in relation to audio logical equipment.
Bidders to perform all relevant pre- and post-acoustic evaluation tests for the soundproof booth according to SANS 10182 Standards when installation complete
Certificates of compliance to be provided
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list