## Sheet1
Vestibular equipment: VNG system
Description
A medical device with a series of tests used, for testing inner ear and central motor functions using infrared goggles to trace eye movements during visual stimulation and positional changes.
Technical specifications
Laptop PC must have one 34mm PCE Express Card slot available (for side mount FireWire camera goggles only)
Desktop PC must have one PCI Express card available (for side mount FireWire camera goggles only)
Must have USB port (expanded by USB hub)
Must have Intel i5 processor 2.5 GHz or better, and not older than 2nd generation
Must have minimum 8 GB RAM or more with windows software
Must have hardrive with minimum of 250 GB space
Must have minimum display of 1366x768 (higher resolution recommended) with a 60Hz sampling rate for all tests.
Must be able to store patient and test data.
Medical electrical equipment needs special precautions regarding electromagnetic compatibility (EMC) and needs to be installed and put into service according to the EMC information provided
Portable RF communications equipment (including peripherals such as antenna cables and external antennas) should be used no closer than 30cm to any part of the VNG including cables specified by the manufacturer.
Use of accessories, transducers, and cables other than those specified or provided by the manufacturer of this equipment could result in increased electromagnetic emissions or decreased electromagnetic immunity of this equipment and result in improper operation.
Contact Micromedical for approved replacement parts
The emissions characteristics of this equipment make it suitable for use in industrial areas and hospitals (CISPR 11 Class A).
If it is used in a residential environment (for which CISPR Class B is normally required) this equipment might not offer adequate protection to radio-frequency communication services.
The user might need to take mitigation measures, such as relocating or re-orienting the equipment
Must have a USB or PC interface
Power supply: input (100-240VAC/50-60Hz/700mA) and output (15V DC/2A)
Must have inputs VNG of at least 4 channels for binocular testing
Must have inputs ENG of 2-4 channels
The CMR ratio must be (>100dB at 50/60Hz)
Must have channel frequency response of at least 12dB/octave low pass filter
Input impedance:
Channel 1 : >5.5 Mohms
Channels 2-4: >8 Mohms
The coupling must have a DC response
The linearity must be 1% full scale horizontal and 1.2% full scale vertical
Must have a sample rate of at least 60Hz for all tests
The eye tracking must be at least 30 degrees horizontal and 30 degrees vertical
Must have a Windows graphical high performance analysis software
Weight:
hardware: not more than 6lbs
light bar : not more than 6lbs
binocular goggles: not more than 450g
PC requirements:
processor: Pentium 4 preferable
RAM: minimum 1G
bus support: USB 2.0
OS: Microsoft XP professional package
CD drive: CD-R/W
display colour: 32bit colour
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list