## Sheet1
A Mobile (portable) Suction/vacuum machine complete for lamination/thermo-moulding
Description
A portable device used for vacuum forming of thermoplastic materials and lamination of prosthetic sockets.
Technical Specifications
The unit must be fully automatic and controlled through 22 contact pressure gauges
The unit to be supplied with 2 separately adjustable vacuum banks, each with no less than 3 connections
It must have water separator with visual and audio controls
Minimum suction power of 25 l/min
The weight of the unit should not exceed 19kg
The unit must have adjustable pressure
The unit provided must be with the following approximately dimensions: 440 x 450 x 230 mm
Vacuum gauge to indicate the vacuum
The device must be supplied with:
replaceable filters,
open-end wrench
10m pvc hose
2 x Water Separator bottles with two hose connection points of 6mmǾ
Accessories:
replaceable filters
open-end wrench
10m pvc hose
water separator bottle with two hose connection points of 6mmǾ
black flow seal 28mmǾ
The unit must operate from 220v with SABS approved 3 pin plug and minimum of 2m in length
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list