## Sheet1
Goniometer Set
Description
An instrument used to measure joint range of motion
Technical Specification
Set should include at least 4 commonly used goniometers in a carrying case with compartments that holds goniometers firmly in place:
A complete set must include the below:
a. Finger goniometer
Used to measure range of motion of the fingers.
Made of durable hard plastic
Must have 180 degrees scale marked in 1 degree increments that read in either direction.
Must be able to measure hyperextension of the finger (0 - 120° with hyperextension 0 - 30°)
Linear in inches and centimeters.
b. Goniometer 15-20cm
Used to measure range of motion of medium joints of the body.
Made of durable hard plastic
Must have 0-180° scale
Must have 1-2° degree increments (Rulong)
Must have linear in inches and centimeters
c. Goniometer 15-20cm
Used to measure range of motion of medium joints of the body.
Made of durable hard plastic
Must have 0-360° scale
Must have 1-2° degree increments
Linear in inches and centimetres
d. Goniometer 30 cm
Used to measure range of motion of large joints of the human body
Made of durable hard plastic
Must have 0-360° scale
Must have 1-2° degree increments
Linear in inches and centimetres
Accessory:
Carry case
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list