## Sheet1
Splint Pan (11-20 litres)
Description
A device used for moulding thermoplastic splinting materials by means of hot water during therapy sessions, generally but not limited to hand therapy.
Technical Specification
Concealed heating element with a perforated tray with non-stick liner to prevent splinting material from adhering to pan base
Should not have water circulated or causing a flow current
have built in syphon drainage
lid with insulated plastic handles for safe removal (if lid is stainless steel)
Must be made of stainless steel
Maximum boiling temparature
Adjustable thermostat controls with thermometer on control panel & indicator light.
Must be durable and rust proof
Electrical cord with SA compatible 3 pin plug, 220 -250 Volt,Electrical cord with SABS approved standard
CE marking standard or equivalent
Approximately size of 11-20 litres
Minimum internal dimensions 500mmx 300mm x 180mm
Accessories:
Spatula
Perforated tray
Non stick liner
For commisioning purposes suppliers must supply 10cm x10cm thermo materials
Optional
Digital
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list