## Sheet1
Parallel Bars 3m, Wall Mounted, Foldable (stainless steel)
Description
A wall mounted, horizontally foldable unit used for gait training (walking) of patients
Technical Specification
Parallel bars designed to be wall mounted and easily folded to the floor and back to the wall, horizontally
Handrails measure at least 3m in length, and fold against the wall vertically when not in use
Folded dimension approximately 25cm deep when folded
Folding mechanism to be ergonomic and durable
Dimensions: ± 275cm(9ft) length x 63cm(25") width x 71cm - 104cm (28" - 41")height
Easily adjustable height of handrails from approximately 55cm – 100cm, in 3-4 cm increments, with telescopic steel tubing with durable locking pins
Fixed width of approximately 63cm
Handrails shall be durable, stainless steel
Handrails shall be 3.8cm in diameter
Heavy gauge steel base plates, uprights and fittings
Each upright telescope up to ±33cm and locks into place
The unit shall have a weight-bearing capacity of 250kg, with sturdy, durable base plates
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list