## Sheet1
Hydrocollator 8 packs, fixed unit
DESCRIPTION
A unit used to warm moist heat packs
Technical Specification
This specification establishes the requirements and supply of a mobile PHYSIOTHERAPY – HYDROCOLLATOR UNIT. The unit offered shall be specifically designed to provide consistent heat for HOTPACKS that will be used to treat patients with conditions requiring moist heat therapy treatment
The unit shall be suitable for a high usage clinic
The hydrocollator heating unit shall have the capacity to accommodate at least 8 heat packs
The unit shall have a durable stainless steel water tank that houses heat packs
The unit shall have a low water sensor that shuts off the unit when the water level gets too low
The unit shall have full insulation to prevent heat loss
The unit offered shall be fitted with a reliable temperature control device to regulate the temperature at a temperature range of between 70º C to 75º C.
The accuracy of the temperature regulation must be within ± 10% or better
The heat up time to 70 º C must not exceed four hours on the unit offered
The cool down time to 70 º C must not exceed two hours on the unit offered
The unit offered must be fitted with a safety temperature cut out device that will shut off power to the heater when it exceeds a temperature in a range of between 82º C to 85º C
The unit offered must incorporate a reliable drain mechanism to drain the water
The unit shall be easy to clean
The unit shall not require any plumbing
The unit shall be easy to maintain. Maintenance specifications to be supplied by bidder
The bidder shall specify the dimensions of the unit
Each unit delivered shall include 6 standard sized (25cm x 30cm) hot packs and 2 cervical hot pack
Each unit delivered shall include heat resistant tongs
Accessories:
castors with locks
heat pack teri covers: standard & cervical (individually priced)
all sizes heat packs
The unit shall operate from 220V and the power cords shall have a 3-point plug approved by SABS
The unit offered must be supplied with internal removable stainless steel racks that will accommodate all sizes and shapes of hot packs
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list