## District Hopital Specifications
PACS and RIS Solution Performance and Usability
1 | Requirements:
1.1 | All individual modules required to provide the functionality as specified in this specification shall be included and licensed to the purchaser.
1.2 | The PACS and RIS solutions provided shall be designed on a high availability (HA) architecture minimizing downtime due to hardware and/or software failures.
1.3 | The PACS and RIS solutions shall be designed to provide a guaranteed uptime of no less than 99% per month, excluding downtime relating to network issues but including downtime due to server hardware and software issues.
1.4 | The PACS and RIS solutions provided shall be designed to provide the same accepted levels of performance as signed off after user acceptance testing (UAT), throughout the 7-year operational use of the solution.
1.5 | PACS performance shall be measured based on the following minimum criteria:
1.6 | The time taken to display the first image of a series from the PACS solution shall not take longer than 3 seconds on any workstation review or diagnostic throughout the hospital.
1.7 | On viewing large volume based CT or MRI studies it shall be possible to scroll through multiple linked series without any jitter or delay between slices no matter how many slices are in the study.
1.8 | On viewing volume based reconstructions like MPR in CT studies transition between views in all planes should be smooth without any jitter or delay between views.
1.9 | The PACS and RIS solutions shall be designed with a user-friendly interface enhancing the overall user experience when using the various applications.
2 | Hardware and Software Infrastructure:Server Invironment
2.1 | Solutions shall be provided for PACS,VR and RIS including server hardware, storage hardware, and software to support the user, modalities types and procedure volumes as noted in the general information section of this specification.
2.2 | The PACS and RIS solutions shall be deployed in a virtual server environment or similar, providing the capacity to enable the performance noted in this specification.
2.3 | Storage capacity for the PACS and RIS solutions shall provide online storage for all types of data for a minimum of 5 years based on the number of procedures and services noted in this specification.
2.4 | The PACS and RIS solutions shall support data redundancy in case of disk failure in the storage component. This shall be configured with a minimum of RAID 5 or better.
2.5 | A disaster recovery solution must be provided. The solution must include any additional hardware,software and professional services to support this environment.The disaster recovery shall allow the recovery of destroyed production environment and all its data, configuration and functionality.
3 | Radiologist Requirements (Institution specific)
3.1 | Diagnostic workstation
3.1.1 | Quantities required will be determined by purchasing institution:
3.2 | General Diagnostic workstation:
3.2.1 | Each workstation shall include 1 X 23” colour monitor and 2 X 3 MP minimum reporting monitor or 1 X 6 MP minimum fusion/dual colour monitor.
3.3 | Mammography Diagnostic workstation:
3.3.1 | Each workstation shall include 1 X 23” colour monitor and 2 X 5 MP minimum reporting monitor or1 X 12 MP minimum fusion/dual colour monitor.
3.3.2 | The diagnostic monitors supplied shall be licensed for import by the Directorate of Radiation Control (RadCon).
4 | Monitor
4.1 | For the Radiologists workstation 1 x 23" Colour LCD minimum monitor shall be supplied.
4.2 | The 23” monitor shall be LCD colour monitor.
4.3 | The size of the monitor shall be 23" or larger
4.4 | The pixel resolution of the 23” monitor shall be at least 1,3 MP
4.5 | The 23” monitor shall be able to have a viewing angle of, at least 15-30º -Please state viewing angle.
4.6 | The refresh rate of the 23” monitor shall be more than 75Hz. Please state refresh rate.
4.7 | All graphics cards, connection cables, and power cables shall be supplied with the monitor.
5 | Medical Grade Monitor
5.1 | Radiologist workstation: 2 x 3 Mega Pixel/ 6 mega pixel (fusion/dual) High Bright Medical monitors shall be supplied according to DICOM part 10 standards.
5.2 | Mammography radiologist workstation: High Bright Medical monitors shall be supplied according to DICOM part 10 standards. For those who don't import these monitors but purchase them from a local supplier, supply proof of the license and a signed confirmation letter from local supplier indicating that they have a distribution and support agreement with the bidder to supply and support the monitors.
5.3 | There should be a concurrent supply of radiologist licenses according to the needs of the institution + 1 floating license.
6 | The following software packages shall be available:
6.1 | MR Spectroscopy – Processing Spectroscopy images
6.2 | DWI – Diffusion weighted imaging (DWI).
6.3 | DTI – Diffusion Tensor Imaging.
6.4 | Orthopedic package (6 Concurrent Licenses).
6.5 | Mammo (tomosynthesis) should be importable from modality:
6.6 | MRI mammo package
6.7 | 3D/4D – Volume Viewer.
6.8 | Neuro - 4D perfusion for neuro and brain view.
6.9 | PET/CT - multi-modality.
6.1 | CT Colonoscopy package.
6.11 | Oncology package.
6.12 | Vascular package.
6.13 | Cardiac package.
6.14 | Maxilo package.
6.15 | Modality workstation integration of CT, MRI and Mammo shall be possible
6.16 | Any other software package might deem applicable but not mentioned above, shall be listed with its funcionality and available options.
6.17 | When reporting it shall be possible to have a drop-down menu of ICD 10 codes and shall be available on each report and approval of report shall not be allowed without ICD10 codes.
6.18 | Smart devices should be able to view images and reports on archive via an android application, Apple IOS application or URL address being connected to hospital network i.e. zero footprint platform for image viewing.
6.19 | The PACS application shall allow the user while being busy with a VR report and performing user specific advanced manipulation to place on hold the report and advanced image manipulation, enter another case perform advanced image manipulation then return back to original case and continue reporting an advanced image manipulation from where he left off. Explain how your solution will achieve this.
6.2 | The PACS system shall allow DICOM images to be exported to an external media like CD, DVD and other external storage media on request by the user.
6.21 | When the DICOM images are exported a DICOM viewing program shall be configured to use the auto playing function.
6.22 | The PACS application on the Radiologist workstation shall allow the user to move export selected images to JPEG format and TIFF format, and be able to adjust the image quality.
6.23 | ICD 10 coding must be included.
6.24 | A digital teaching file system option must be incorporated into the PACS.
7 | Image Manupulation Tools
7.1 | Flipping images on the Horizontal and vertical planes shall be possible on all Radiologist workstations.
7.2 | Rotating images 90 and 180 degrees shall be possible on all Radiologist workstations.
7.3 | Pre programmed window width and level settings shall be possible on all Radiologists workstations.
7.4 | Dynamic interactive brightness and contrast control selected on the mouse shall be possible on all Radiologists workstations.
7.5 | Dynamic interactive zoom in and out directly selected on the mouse shall be possible on all Radiologists workstations.
7.6 | Dynamic interactive PAN feature shall be possible on all Radiologists workstations.
7.7 | The user shall be able to activate a cine loop functionality, and customize the speed.
7.8 | The Radiologist PACS application work list shalldisplay the overall Patient examination history once display the overall Patient examination history once a patient has been selected.
7.9 | The user shall be able to set the speed of the cine loop to correspond to the actual cine speed when the data was acquired. This shall be the default setting.
7.1 | The user shall be able to customize whether the cine loop automatically repeats itself.
7.11 | The user shall be able to customize the direction of the repeat. E.g. begin to end & end to begin OR begin to end & begin to end.
7.12 | The user shall be able to link the cine functionality (described above) across two or more frames.
7.13 | The solution shall intelligently match the correct slice locations of the selected series/study. i.e. if 2 studies are being compared the slice location in each study shall be the same or as close as possible.
7.14 | The user shall be able to manually link two or more series/studies. This functionality will be used when the automatic link is inaccurate and will override the automatic link is inaccurate and will override the automatic link.
7.15 | The user shall be able to link image manipulation functionality across multiple frames. The linkable functionality shall be:
7.15.1 | Magnifications.
7.15.2 | Pan.
7.15.3 | Window width and level.
7.15.4 | It shall be possible for the user to apply imaging enhencement filters.
7.15.5 | Edge enhancement shall be possible
7.15.6 | Image sharpening shall be possible
7.15.7 | Image softening shall be possible.
7.15.8 | Scrolling in multiple plains
7.15.9 | Reference lines to show relative position.
8 | Measurement tools
8.1 | Radiologists PACS application shall be capable of measuring distance on the displayed image.
8.2 | Multiple distance measurements on a single image shall be possible on all Radiologist workstations.
8.3 | The Radiologists PACS application shall be capable of measuring an angle on the displayed image
8.4 | Multiple angle measurements on a single image shall be possible on all Radiologist workstations.
8.5 | The basic imaging manipulation interface of the imaging solution shall be capable of measuring a COBB angle on the displayed image.
8.6 | Multiple COBB angle measurements on a single image shall be possible on all Radiologist workstations.
8.7 | The Radiologist PACS application shall be capable of measuring a region of interest (ROI) on the displayed image.
8.8 | The measurement shall reflect average, minimum,maximum and standard deviation of the density within the ROI.
8.9 | Multiple ROI measurements on a single image shall be possible on all Radiologist workstations.
8.1 | Shapes shall be available for the ROI measurement on each Radiologist workstation. Circular, Rectangular, freehand drawn with mouse
8.11 | The reference points for any measurements shall be editable by the user. It shall be possible for the user to make text annotations
8.12 | It shall be possible to delete measurements.
8.13 | Any image to which text or measurements have been added will automatically be stored as a key image.
9 | Quality Assurance
9.1 | Shall conform with the latest Mammography Quality Standards Act (MQSA) regulations
9.2 | Quality assurance calibration shall be built into monitors to perform tasks such as luminance measurement and other quality assurance tests that are needed from radiation control for diagnostic monitors. If calibration tools are not build into monitors then calibration tools shall be included with diagnostic monitors.
9.3 | AAMP TG 18 test patterns to be included with PACS and shall always be available on cache storage, it shall also be able to view dually on both diagnostic monitors. Explain your solution.
9.4 | L-angle distance patterns shall be included with PACS and shall aslo be available on on-line storage.
10 | Worklist Management
10.1 | The Radiologist PACS application shall support configurable work list management.
10.2 | The Radiologist PACS application shall have the contingency measure to search the PACS database for patient's Images, outside of the RIS application.
10.3 | The PACS application shall allow the user to search on the following DICOM fields from within the PACS application:
10.3.1 | Patients Last Name
10.3.2 | Patient Unique ID number or (MRN)
10.3.3 | Study Date
10.3.4 | Accession Number
10.3.5 | Study Description
10.3.6 | Referring Physician
10.3.7 | Modality
11 | The Radiologist PACS application shall allow for pattern searching within each of these fields. (Pattern searching refers to searching for a specific pattern of text or numerals, or symbols within the field).
12 | The Radiologist PACS application shall allow the User to use wildcards within the search criteria.
13 | The presentation of the work list to the user shall be configurable.
14 | The columns that are presented to the user should be configurable based on the DICOM tag required. Please state what fields are configurable.
15 | The columns that are presented to the user shall be able, to be sorted ascending or descending.
16 | The order in which the columns are displayed should be configurable.
17 | The Radiologist PACS application shall open the appropriate patient images on the viewing monitors after the patient has been selected from the work list.
18 | The Radiologist PACS application work list shall display the overall Patient examination history once a patient has been selected.
19 | Hanging Protocols
19.1 | When using the basic image manipulation app on a multi-monitor workstation, the user will be able to use hanging protocols. When using a single-monitor display it will be displayed according to the last and highest ranking presentation state, saved by a previous user.
19.2 | The Hanging Protocols shall be configurable.
19.3 | User defined hanging protocols shall be available to A specific users
19.4 | System wide define hanging protocols shall be available to all users.
19.5 | The configurations of the Hanging protocols shall include the ability to change the screen layouts to user define layouts.
19.6 | It shall be possible to change the layouts per monitor.
19.7 | It shall be possible to change the layouts across monitors. Hanging protocols must load selected image display features automatically, for instance-reference lines, synchronisation, zoom and pan, mirror image.
19.8 | It shall be possible to configure what is displayed in each segment of the display layout selected. Hanging protocols must load the advanced image manipulation application, and it must be a multi-modality hanging protocol.
19.9 | Which images are hung in which segment shall be configurable based on the following criteria:
19.9.1 | Study date
19.9.2 | Study time
19.9.3 | Body part
19.9.4 | The presence of previous patient
19.9.5 | examinations
19.9.6 | Modality
19.9.7 | Windows level
19.9.8 | Examination type
19.9.9 | Presentation states
19.9.10 | Prior vs curent
19.9.11 | Series description
19.9.12 | Study description
19.9.13 | Body position
19.9.14 | Laterality
19.1 | The Radiologist PACS application shall allow the User to save and label presentation states.
19.11 | These presentation states shall be saved as a presentation state object as part of the DICOM record in the PACS system
19.12 | The user shall be able to change the font size and colour of the annotations.
19.13 | The user shall be able to add "arrow" annotations to the image.
19.14 | Multiple annotations shall be possible on a single image.
19.15 | A pre defined list of annotations shall be available for the user.
19.16 | The user shall be able to edit the positions of an annotation.
19.17 | It shall be possible to show/hide annotations
19.18 | Spine annotations shall be possible in which the annotation automatically changes based on the previous annotation .e.g. L1 automatically changes to L2.
19.19 | Please state any other specific annotation functions your software solution provides.
19.2 | Annotations shall be saved as an overlay on the image.
20 | DICOM PRINTING
20.1 | Authorized users may print to a DICOM printer from the basic image manipulation application.
20.2 | The Radiologist PACS application DICOM print interface shall allow the user to select customizable layout formats.
20.3 | Please list the available print layouts your PACS application supports.
20.4 | The Radiologist PACS application DICOM print interface shall allow the user to Select different film sizes.
21 | Key Images
21.1 | Any images with annotations and measurements will automatically be saved as key images.
21.2 | These key images shall be saved as key images objects as part of the DICOM record in the PACS system.
22 | Desktop Intergration BETWEEN RIS AND PACS (Institution Depended)
22.1 | The Radiologist PACS applications shall integrate seamlessly with the RIS application. A RIS/PACS solution is required from the vendor with seamless integration to HIS. If HIS is not avaialbale at institution, provision will be made for integration once HIS is deployed. i.e the institution shall be future proof for HIS integration.
22.2 | This integration shall automatically display the correct patient images in the PACS application when the patient is selected in the RIS. Please state how this is done.
23 | MULTIPLANAR RECONSTRUCTION (MPR)
23.1 | It shall be possible to construct a MPR view of CT or MRI studies from within the Radiologists PACS software.
23.2 | The user shall be able to see at least three frames
23.3 | The three frames shall be populated by a sagittal, coronal, and axial view of the loaded study.
23.4 | Positional correction – Each plane is defined by its relative position in two other planes. E.g. an axial plane is defined by its relative position in both the coronal (to define the left to right angulations) and the sagittal plane (to define the anterior to posterior angulation).
23.5 | The PACS application shall show two orthogonal lines per plane, demonstrating the relative position of the other two planes on the current plane i.e. reference lines.
23.6 | The user shall be able to view the orthogonal lines on all planes at all times.
23.7 | The user shall be able to hide/show the orthogonal lines.
23.8 | The user shall have an indication of which Orthogonal line is referencing which plain. The angle between the plains shall always be 90 Degrees in MPR mode.
23.9 | It shall be possible for the user to change the Position of the orthogonal lines dynamically.
23.1 | When an orthogonal line is changed the corresponding plane shall update in real time.
23.11 | It shall be possible to create a slab of data in any plane.
23.12 | The thickness of the volume shall be able to be dynamically adjusted by the user.
23.14 | The user shall be able to dynamically adjust the thickness of the volume.
23.15 | The user shall be able to choose between the following different intensity projections of the data within the volume.
23.16 | MIP (Maximum Intensity Projection) shall be possible.
23.17 | AveIP (Average Intensity Projection) shall be possible.
23.18 | MinIP (Minimum Intensity Projection) shall be possible
24 | Other Multiplanar Reconstruction Modes
24.1 | Multiplanar Reconstruction (MPR) Double Oblique
24.1.1 | It shall be possible for the user to perform MPR Double Oblique function within the Radiologist PACS Application.
24.1.2 | The only difference between the MPR Double Oblique function and the standard MPR function shall be that the angle between the planes can change.
24.1.3 | All other function describe in the standard MPR requirements shall be possible in the Double Oblique modes as well.
24.1.4 | Please state if any functions are not possible.
24.2 | Multiplanar Reconstruction (MPR) Curved Oblique
24.2.1 | It shall be possible for the user to preform a MPR
24.2.2 | Curved Oblique function within the Radiologist PACS Application.
24.2.3 | In the curved MPR Mode the user shall be able to see at least four frames.
24.2.4 | The frames shall contain the sagital, coronal, and axial planes. This will be used to place the points to create the curved oblique plane.
24.2.5 | The fourth frame shall contain the curved oblique plane.
24.2.6 | The user shall be able to mark points on any of the planes.
24.2.7 | The line connecting the points shall be created automatically.
24.2.8 | The user shall be able to edit the position of any of the points already displayed.
24.2.9 | The user shall be able to rotate the curved MPR using the curved oblique line as the axis of rotation
24.2.10 | It shall be possible to save any MPR from any mode to the PACS application for future review.
25 | Voice Recognition
25.1 | The voice recognition shall seamlessly integrate with the PACS and RIS.
25.2 | The voice recognition shall be set up for each individual user/radiologist to enable the system to learn from each unique speech of the users/radiologists.
25.3 | When a radiologist is no longer present at the healthcare facility, it shall be possible to revoke his speech license and make it available for a new user.
25.4 | There shall be one floating license that can be used by more than one user.
25.5 | The solution shall be able to convert Conversational speech into the highest accurate written text on a PACS report.
25.6 | There shall be no need for restrictive voice commands for punctuation or formatting.
25.7 | The solution shall have an Automatic and continuous learning system.
25.8 | The voice recognition shall be able to adjust to a variety of indigenous accents to ensure best results.
26 | RIS Functional Requirements
26.1 | The patient banner information to be displayed on the top demographic banner of any clinical system (PACS, RIS) should include the following attributes: name, DOB, sex, MRN number, ID Number, current patient location and current responsible consultant. Internal Pseudo-MPI (Master Patient Index) capabilities to tie individual facility MRN’s to the same patient in a multi-facility organization.
26.2 | The status should be synchronized between RIS and PACS and should consist of Arrived, Exam Started, Exam Completed, Exam not performed with reason, Verified, Failed verification and Finished Exam Explain if this is possible with your solution.The solution offered shall be able to automate the start and finish of exam statuses.
26.3 | The DICOM Modality Work list shall be available on all the modalities at the institution. Once the status of the exam changes to “EXAM PERFORMED” or “EXAM NOT PERFORMED” on RIS, the exam shall be removed from the work list automatically.
26.4 | The following information needs to be provided by RIS to modalities: Patient Demographics: Name, DOB, Sex, Modality, Exam Description, Exam Code, Exam Room, Accession No, MRN number, ID Number.
26.5 | Three click reporting workflow shall be possible from the RIS & PACS integration as follows: Draw up a RIS worklist for reporting; Launch a patient’s exam for reporting on RIS - automatic display of images on PACS with automatic display of relevant prior - dictate/VR a report on RIS; Click on the next exam on RIS worklist - which closes the previous PACS images & launches the next patient on RIS & PACS.
26.6 | Preliminary reports shall be available on system and shall be able to move from preliminary status to approved status by a click of a button.
26.7 | Previous images of a study shall be available for comparison purposes and shall be displayed simultaneously on left and right monitors without it being necessary to close windows.
26.8 | All approved reports shall be available on HIS (if institution has HIS). Explain if this is possible with proposed solution.
26.9 | A drop-down menu of ICD 10 codes shall be available on each report and approval of report shall not be allowed without ICD 10 code.
26.1 | Each Radiologists User shall have access to a personal folder where special patient can be stored.
26.11 | Final report shall be simultaneously displayed with its relevant images and if an image has a final report then the report shall be displayed simultaneously.
27 | PACS Functional Requirements
27.1 | User access control, security and audit tracking
27.1.1 | The access to the PACS solution shall be controlled by user access control, security and user tracking. These functions shall be based on sound IT principles. The following section outlines the requirements for RIS access, security and tracking.
27.1.2 | The PACS solution shall restrict access to only authorised users.
27.1.3 | The PACS solution shall support individual user login with password authentication.
27.1.4 | The PACS solution shall support role-based access controls to restrict access to system functionality based on predefined user group privileges.
27.1.5 | The PACS solution shall provide an audit trail of all user actions performed within the PACS solution, these audit trails shall be accessible to authorised PACS users.
27.1.6 | The PACS solution shall provide an audit trail of all user access to and modification of patient data, these audit trails shall be accessible to authorised PACS users.
27.1.7 | It shall be possible to mark individual patients records as confidential and restrict access to a specific user group.
27.2 | Data Management
27.2.1 | Data management is the functionality in the PACS which allows an authorised user to manage the data lifecycle of data stored in the archive. This includes moving data from expensive storage hardware to less expensive storage hardware based on a set of data lifecycle rules. It also includes managing retention periods of studies and when and which studies can be deleted.
27.2.2 | The PACS shall support rule-based lifecycle data management, to manage the movement of data in a tiered archive solution.
27.2.3 | The PACS shall support rule-based lifecycle data management, to manage the retention of data in the archive solution. Retention of data is achieved by not allowing users the function to delete studies if the study is protected by the retention rule.
27.2.4 | The PACS solution shall support automatic rule-based routeing of data to different nodes.
27.2.5 | The basic PACS functionality is functionality that is traditionally provided by the PACS solution and applicable to all authorised users. The specifications for base PACS functionality are detailed below.
27.2.6 | It shall be possible to change window level and centre.
27.2.7 | It shall be possible to store pre-set window settings for CT viewing.
27.2.8 | It shall be possible to change the image layout on the viewing screen.
27.2.9 | It shall be possible to measure distance, angle, Cobb angle, a Region of interest, Hounsfield units, ext.
27.2.10 | It shall be possible to manually scroll through large cross-sectional studies like CT, Angiography and MRI.
27.2.11 | It shall be possible to automatically scroll through large cross-sectional studies with a cine loop function.
27.2.12 | It shall be possible to vary the speed and direction of the cine loop.
27.2.13 | It shall be possible to link multiple series in a cross-sectional study like CT and MRI and scroll through the series simultaneously.
27.2.14 | It shall be possible for an authorised user to create custom shortcut keys for commonly used functions.
27.2.15 | It shall be possible for an authorised user to export images as jpeg images.
27.2.16 | It shall be possible for an authorised user to export images as Tiff images.
27.3 | Radiographers PACS Functionality
27.3.1 | The Radiographer PACS functionality refers to PACS functionality that is required, which is specifically related to Radiographers. This functionality should be available to authorised radiography users working on any workstations in the department, especially the RIS workstations which are located next to or in the region of modalities. The specifications for PACS features relating to radiographers are detailed below.
27.3.2 | The PACS shall provide for pre-configured text labelling
27.3.3 | The PACS shall allow authorised users to label and save presentation states.
27.3.4 | The PACS shall allow authorised users to print DICOM films to the DICOM laser printer.
27.3.5 | It shall be possible for an authorised user to send DICOM studies to the CD/DVD robot. (Institution Optional)
27.4 | Radiologists PACS Functionality
27.4.1 | The Radiologists PACS functionality refers to functionality that is required to view and diagnose radiology images in the radiology department. This functionality should be available to radiologists working on dedicated diagnostic workstations in the department and at home when they are on call. The specifications for PACS features relating to radiologists are detailed below.
27.4.2 | The PACS shall allow authorised users to create user-specific function keys for commonly used functions.
27.4.3 | The PACS shall allow authorised users to select and store key images.
27.4.4 | Key images stored in the PACS shall be stored as DICOM key image objects.
27.4.5 | Annotations, presentation overlays, presentation states stored as part of the DICOM study shall be visible to the authorised user.
27.4.6 | The PACS shall allow authorised users to add annotations, label and save presentation states.
27.4.7 | Presentation states saved to the PACS shall be stored as DICOM presentation state objects.
27.4.8 | The PACS shall allow authorised users to configure user-specific hanging protocols.
27.4.9 | The solution shall provide the authorised users with a list of recently viewed studies.
27.4.10 | It shall be possible for authorised users to flag interesting cases and save them in a teaching file folder.
27.4.11 | It shall be possible for authorised users to store interesting cases.
27.4.12 | The solution shall automatically link 2, 3 or 4 series within a cross-sectional type study based on a predefined hanging protocol and synchronise scroll through the images.
27.4.13 | The solution shall allow an authorised user to manually link 2, 3 or 4 series within a cross-sectional type study and synchronise scroll through the images.
27.4.14 | The solution shall provide for remote reporting services by allowing a user to login to the system and access functionality remotely. This will include creating reports and accessing images.
27.5 | PACS Advanced Functionality
27.5.1 | The PACS solution shall provide basic advanced visualization tool applicable to volumetric studies as specified below:
27.5.2 | The PACS functionality shall support Multi-Planar reconstructions (MPR) in the following planes, axial, coronal, sagittal, and oblique.
27.5.3 | The PACS functionality shall support Curved-Planar reconstructions (CPR) including projected, straightened and stretched views.
27.5.4 | The PACS functionality shall support changing slab slice thickness.
27.5.5 | The PACS functionality shall support slab rendering based on maximum (MIP), minimum (MinIP) and average (AvgIP) intensity projections.
27.5.6 | The PACS functionality shall support 3D volume rendering for specific anatomy parts such as lung, bone, vessels and skin.
27.5.7 | The PACS functionality shall support rendering 3D animations and exporting these animations as AVI files.
27.5.8 | The PACS functionality shall support basic segmentation tools, including bone removal, cropping, clipping, freehand selection.
27.5.9 | The PACS functionality shall support using selected advanced visualization tools automatically in predefined hanging protocols.
27.5.10 | The PACS functionality shall support volume measurements.
27.5.11 | The PACS shall support basic dental specific functionality, please give a detailed description of your solutions specific dental functionality. (Institution Specific)
27.5.12 | The PACS must have compatability for possible future Artificial Intelligence (AI) applications.
28 | Clinicians PACS Functionality
28.1 | One of the critical functions the solution shall support is a clinician portal. This is the portal that the clinicians will use to access order scheduling information, images and reports. The DoH is looking to take advantage of the new mobile technologies and zero footprint viewers throuhg a web based clinicians portal
28.2 | The solution shall support viewing of order information, images and reports through a web-based clinicians portal.
28.3 | It shall be possible for authorised users to search for patients based on their Patient Master Index identifier, Patient name, study date, and/or study date range
28.4 | In the case of reports created by registrars and signed off by a consultant, both the preliminary and final report shall be available.
28.5 | The solution shall support basic PACS functionality as noted through the clinicians portal.
28.6 | The solution shall support basic visualisation tools like MPR, MIP and 3D rendering through the clinicians portal.
28.7 | The solution shall support access to the portal via mobile technology such as smartphones, and tablets or on a browser through URL.
28.8 | The solution shall support cross-browser viewing.
28.9 | The solution shall support Android devices.
28.1 | The solution shall support Apple devices.
28.11 | The solution shall support Windows devices.
28.12 | The solution shall support access to images through an HTML5 web browser.
28.13 | The data streamed from the PACS or RIS archive to the device shall be based on a secure encrypted service.
28.14 | No image or report data shall remain stored on the device once the connection is dropped or disconnected.
28.15 | For the transcriptionist provision shall be made for:
28.15.1 | Transcriptionist Workstations (Institution optional)
28.15.2 | USB transcriptionist set including
28.15.3 | headset
28.15.4 | and paddle (Institution optional)
28.15.5 | Provision shall be made based on an institutional requirements
28.15.6 | RIS Workstations
28.15.7 | PACS Administrator Workstation and
28.15.8 | or Notebook and or Tablet
28.15.9 | Orthopeadic Workstation
28.15.10 | Radiogrpaher Viewing Workstation
28.15.11 | All Network equipment shall be able to connect 1Gbs Network.
28.15.12 | All computer equipment shall include keyboard and mouse.
28.16 | Should the institution not have the necessary network requirements for the system to operate optimally, a network quotation will be supplied.
28.17 | The following should be considered when developing the network quotation:
28.17.1 | The network shall have a fiber uplink to the core switch located in the server room at a minimum of 1Gb/s to connect to DoH network.
28.17.2 | Layer 3 manageable switch is required with PoE with redundant/spare switch available.
28.17.3 | Current cabinets may be utilized and all newly installed cabling must be marked clearly.
28.18 | A site survey is recommended before the proposal is submitted to ensure the correct equipment is installed.
29 | The final proposal must be approved by End Users
29.1 | Vendors shall provide a full diagram and description of their whole solution including RIS-PACS, integration with HIS (institution optional), data migration,archiving options, server details,primary Archive The PACS shall be able to send images to other institutions and to a future central data repository or VNA, it shall also be able to receive images from other institutions i.e. the PACS shall include the DICOM Service Class license Query/Retrieve SCU and SCP.
29.2 | The solution offered shall make provision for 5 years online (cached) storage.
29.3 | In the event of a failure within the system, the system will automatically alert the PACS administrator and vendor of potential problems. A Total pro-active self-monitoring system of server room shall be provided and alerts shall also be done in the form of sms or any other messaging system to alert PACS administrator and vendor. Monitoring shall form part of contract.
29.4 | A full Disaster Recovery solution shall be provided. This implies that a separate archive solution shall be provided in a separate location within the hospital grounds or elsewhere. In the case of a complete disaster of the primary archive the disaster recovery solution shall provide the radiology department with a working archive solution seamlessly. This solution shall provide the same functionality as the primary archive.
29.5 | There should be a responsibity to liaise with modality suppliers and ensure correct connection of all modalities to the PACS server. All modalities shall make useof modality worklist.
29.6 | There should be a smooth transition from hospital power supply to UPS power when there is power-outages. Switch over configuration shall be seamless to switch over and uninterrupted. UPS for RIS/PACS servers to be responsibility of vendor. UPS shall be able to maintain servers for minimum an hour.
29.7 | Own UPS’s should be supplied - independent of the Departmental UPS’s currently available
29.8 | There shall be 2 X aircons per server room for the main and secondary server rooms, where 1 aircon will be used and the other will be off asa redundant measure. Aircons shall be maintaned.
30 | Interoperability Requirements
30.1 | DICOM
30.1.1 | Provide DICOM conformance statement for the PACS solution offered.
30.1.2 | Provide DICOM conformance statement for the RIS solution offered.
30.1.3 | Provide the DICOM conformance statement for the CD/DVD Robot solution. (Institution Specific)
30.1.4 | The PACS solution shall support the DICOM storage service as SCP and SCU.
30.1.5 | The PACS solution shall support the DICOM query and retrieve service as SCP and SCU.
30.1.6 | The PACS solution shall support the DICOM print service as SCU.
30.1.7 | The PACS solution shall support the DICOM Modality Performed Procedure Step (MPPS) service as SCP. P
30.1.8 | The PACS solution shall support the DICOM Storage commitment service as SCP.
30.1.9 | The solution shall the DICOM Modality Worklist Service as SCP. This service may be implemented in the PACS or RIS solution. Indicate which solution will be providing this service.
30.1.10 | The solution shall support all required DICOM SOP classes to store DICOM objects based on the modalities listed herein.
30.2 | IHE
30.2.1 | The solution shall support the IHE Patient Demographics Query (PDQ) profile.
30.2.2 | The solution shall support the IHE Patient Identifier Cross-Referencing (PIX) profile.
30.2.3 | The solution shall support the IHE Consistent Time (CT) profile.
30.2.4 | The solution shall support the IHE Scheduled Workflow (SWF) profile.
30.2.5 | The solution shall support the IHE Patient Information Reconciliation (PIR) profile.
30.2.6 | The solution shall support the IHE Consistent Presentation of Images (CPI) profile.
30.2.7 | The solution shall support the IHE Audit Trail and Node Authentication (ATNA) profile with special reference to the Radiology defined options.
30.2.8 | Required to submit IHE integration statements for all IHE profiles which the solution supports. This should include IHE integration statements for the PACS, the RIS and the CD Robot solutions.
30.3 | HL7
30.3.1 | Provide an original OEM HL7 integration document for the PACS solution which documents which HL7 message types, data types, and data fields are supported by your PACS solution.
30.3.2 | Provide an original OEM HL7 integration document for the RIS solution which documents which HL7 message types, data types, and data fields are supported by your RIS solution.
30.3.3 | The solution shall support an inbound and outbound HL7 messaging interface.
30.3.4 | The solution shall support the following HL7 message types:
30.3.4.1 | ADT- Admission Discharge Transfer
30.3.4.2 | Provide evidence of requirement by referencing appropriate sections in your HL7 integration documents.
30.3.5 | The solution shall support the following HL7 message type:
30.3.5.1 | ORM- Order Management
30.3.5.2 | Provide evidence of requirement by referencing appropriate sections in your HL7 integration documents.
30.3.6 | The solution shall support the following HL7 message type:
30.3.6.1 | ORR- Order Response (to any ORM)
30.3.6.2 | Provide evidence of requirement by referencing appropriate sections in your HL7 integration documents.
30.3.7 | The solution shall support the following HL7 message type:
30.3.7.1 | ORU- Observation Report Unsolicited
30.3.7.2 | Provide evidence of requirement by referencing appropriate sections in your HL7 integration documents.
30.3.8 | Provide a list of HL7 trigger points within the solution. A trigger point is a point at which the RIS or PACS will trigger an outbound HL7 message.
30.3.9 | Data Mining- the system must be able to show studies statistics for all exams performed.