## PACS Tertiary Hospital
PACS Tertiary Specifications
1 | General Requirements
1.1 | The solution for both RIS and PACS shall be based on an unlimited user’s license model for all types of users and advanced software base licenses. Radiologist, Radiographers, Clerks and system Administrator licenses must be made available.Clinicians license shall be unlimited.
1.2 | All individual modules required to provide the functionality as specified in this specification shall be included and licensed to the purchaser.
1.3 | All operating system Licenses for the PACS, RIS and VR solutions shall be included and licensed to the purchaser
1.4 | All database Licenses for the PACS and RIS databases shall be included and licensed to the purchaser.
1.5 | All virtualization software licenses required to support the PACS,VR and RIS solutions shall be included and licensed to the purchaser.
2. | PACS, RIS and VR Solution Performance and Usability
2.1 | The PACS, RIS and VR solutions provided shall be designed on a High Availability (HA) architecture minimizing downtime due to hardware and/or software failures.
2.2 | All individual modules required to provide the functionality as specified in this specification shall be included and licensed to the purchaser.
2.3 | The time taken to display the first image of a series from the PACS solution shall not take longer than 3 seconds on any workstation, be it review or diagnostic, throughout the hospital.
2.4 | On viewing large volume based CT or MRI studies it shall be possible to scroll through multiple linked series without any jitter or delay between slices no matter how many slices are in the study.
2.5 | On viewing volume-based reconstructions like MPR in CT studies transition between views in all planes should be smooth without any jitter or delay between views.
2.6 | The PACS, RIS and VR solutions shall be designed to provide a guaranteed uptime of no less than 99% per month, excluding downtime relating to network issues but including downtime due to server hardware and software issues.
2.7 | The PACS, RIS and VR solutions provided shall be designed to provide the same accepted levels of performance as signed off after User Acceptance testing (UAT), throughout the 7-year (2 years warranty and 5 years maitenance) operational use of the solution.
2.8 | PACS performance shall be measured based on the following minimum criteria:
2.9 | The PACS, RIS and VR solutions shall be designed with a user-friendly interface enhancing the overall user experience when using the various applications.
2.10 | The Voice Recognition licenses shall be on a floating concurrent user licenses and the bidder is required to describe the voice recognition method license methodand the advantages of the method.
2.11 | Concurrent Voice Recognition user licenses must be supplied with all modalities.The number of user licences depend on the buying institution setup.
2.12 | Describe licensing model for voice recognition licensesand how it support multiple users with concurrent licenses.
2.13 | The Voice Recognition licenses must be transferable between users in the case of a user leaving the institution.
3 | Hardware and Software Infrastructure :Server Environments
3.1 | Provide solutions for PACS, RIS and VR including server hardware, storage hardware, and software to support the user, modalities types and procedure volumes as noted in the general information section of this specification.
3.2 | A disaster recovery solution must be provided. The solution must include any additional hardware,software and professional services to support this environment.The disaster recovery shall allow the recovery of destroyed production environment and all its data, configuration and functionality.
3.3 | The PACS, RIS and VR solutions shall be deployed in a virtual server environment or similar, providing the capacity to enable the performance noted in this specification.
3.4 | Storage capacity for the PACS, RIS and VR solutions shall provide online storage for all types of data for a minimum of 5 years based on the number of procedures and services noted in this specification.
3.5 | The PACS, RIS and VR solutions shall support data redundancy in case of disk failure in the storage component. This shall be configured with a minimum of RAID 5 or better.
4. | RADIOLOGISTS REQUIREMENTS (Institution specific)
4.1 | Diagnostic workstation
4.1.1 | Quantities required will be determined by purchasing institution.
4.1.2 | General Diagnostic workstation: Each workstation shall include 1 X 23” colour monitor and 2 X 3 MP minimum reporting monitor or 1 X 6 MP minimum fusion/dual colour monitor
4.1.3 | Mammography Diagnostic workstation: Each workstation shall include 1 X 23” colour monitor and 2 X 5 MP minimum reporting monitor or1 X 12 MP minimum fusion/dual colour monitor.
4.1.4 | The diagnostic monitors supplied shall be licensed for import by the Directorate of Radiation Control (RadCon).
4.2 | Monitor
4.2.1 | For the Radiologists workstation 1 x 23" Colour LCD minimum monitor shall be supplied.
4.2.2 | The 23” monitor shall be LCD colour monitor.
4.2.3 | The size of the monitor shall be 23" or larger
4.2.4 | The pixel resolution of the 23” monitor shall be at least 1,3 MP
4.2.5 | The 23” monitor shall be able to have a viewing angle of, at least 15-30º -Please state viewing angle.
4.2.6 | The refresh rate of the 23” monitor shall be more than 75Hz. Please state refresh rate.
4.2.7 | All graphics cards, connection cables, and power cables shall be supplied with the monitor.
4.3 | Medical Grade Monitor
4.3.1 | Radiologist workstation: 2 x 3 Mega Pixel/ 6 mega pixel (fusion/dual) High Bright Medical monitors shall be supplied according to DICOM part 10 standards.
4.3.2 | Mammography radiologist workstation: High Bright Medical monitors shall be supplied according to DICOM part 10 standards. Who don't import these monitors but purchase them from a local supplier, supply proof of the license and a signed confirmation letter from local supplier indicating that they have a distribution and support agreement with the bidder to supply and support the monitors.
4.3.3 | The vendor shall supply concurrent radiologist licenses according to the needs of the institution + 1 floating license.
4.4 | The following software packages shall be available:
4.4.1 | MR Spectroscopy – Processing Spectroscopy images
4.4.2 | DWI – Diffusion weighted imaging (DWI).
4.4.3 | DTI – Diffusion Tensor Imaging.
4.4.4 | Orthopedic package (6 Concurrent Licenses).
4.4.5 | Mammo (tomosynthesis) should be importable from modality:
4.4.6 | MRI mammo package
4.4.7 | 3D/4D – Volume Viewer.
4.4.8 | Neuro - 4D perfusion for neuro and brain view.
4.4.9 | PET/CT - multi-modality.
4.4.10 | CT Colonoscopy package.
4.4.11 | Oncology package.
4.4.12 | Vascular package.
4.4.13 | Cardiac package.
4.4.14 | Maxilo package.
4.4.15 | Modality workstation integration of CT, MRI and Mammo shall be possible
4.4.16 | Any other software package which might deem applicable but not mentioned above, shall be listed with its funcionality and available options.
4.4.17 | When reporting it shall be possible to have a drop-down menu of ICD 10 codes and shall be available on each report and approval of report shall not be allowed without ICD10 codes.
4.4.18 | Smart devices should be able to view images and reports on archive via an android application, Apple IOS application or URL address being connected to hospital network i.e. zero footprint platform for image viewing.
4.4.19 | The PACS application shall allow the user while being busy with a VR report and performing user specific advanced manipulation to place on hold the report and advanced image manipulation, enter another case perform advanced image manipulation then return back to original case and continue reporting an advanced image manipulation from where he left off.
4.4.20 | The PACS system shall allow DICOM images to be exported to an external media like CD, DVD and other external storage media on request by the user.
4.4.21 | When the DICOM images are exported a DICOM viewing program shall be configured to use the auto playing function.
4.4.22 | The PACS application on the Radiologist workstation shall allow the user to move export selected images to JPEG format and TIFF format, and be able to adjust the image quality.
4.4.23 | ICD 10 coding must be included.
4.4.24 | A digital teaching file system option must be incorporated into the PACS.
4.5 | Image Manupulation Tools
4.5.1 | Flipping images on the Horizontal and vertical planes shall be possible on all Radiologist workstations.
4.5.2 | Rotating images 90 and 180 degrees shall be possible on all Radiologist workstations.
4.5.3 | Pre programmed window width and level settings shall be possible on all Radiologists workstations.
4.5.4 | Dynamic interactive brightness and contrast control selected on the mouse shall be possible on all Radiologists workstations.
4.5.5 | Dynamic interactive zoom in and out directly selected on the mouse shall be possible on all Radiologists workstations.
4.5.6 | Dynamic interactive PAN feature shall be possible on all Radiologists workstations.
4.5.7 | The user shall be able to activate a cine loop functionality, and customize the speed.
4.5.8 | The Radiologist PACS application work list shalldisplay the overall Patient examination history once display the overall Patient examination history once a patient has been selected.
4.5.9 | The user shall be able to set the speed of the cine loop to correspond to the actual cine speed when the data was acquired. This shall be the default setting.
4.5.10 | The user shall be able to customize whether the cine loop automatically repeats itself.
4.5.11 | The user shall be able to customize the direction of the repeat. E.g. begin to end & end to begin OR begin to end & begin to end.
4.5.12 | The user shall be able to link the cine functionality (described above) across two or more frames.
4.5.13 | The solution shall intelligently match the correct slice locations of the selected series/study. i.e. if 2 studies are being compared the slice location in each study shall be the same or as close as possible.
4.5.14 | The user shall be able to manually link two or more series/studies. This functionality will be used when the automatic link is inaccurate and will override the the automatic link is inaccurate and will override the automatic link.
4.5.15 | The user shall be able to link image manipulation functionality across multiple frames. The linkable functionality shall be:
4.5.16 | Magnifications.
4.5.17 | Pan.
4.5.18 | Window width and level.
4.5.19 | It shall be possible for the user to apply imaging enhencement filters.
4.5.20 | Edge enhancement shall be possible
4.5.21 | Image sharpening shall be possible
4.5.22 | Image softening shall be possible.
4.5.23 | Scrolling in multiple plains
4.5.24 | Reference lines to show relative position.
4.6 | Measurement tools
4.6.1 | Radiologists PACS application shall be capable of measuring distance on the displayed image.
4.6.2 | Multiple distance measurements on a single image shall be possible on all Radiologist workstations.
4.6.3 | The Radiologists PACS application shall be capable of measuring an angle on the displayed image
4.6.4 | Multiple angle measurements on a single image Shall be possible on all Radiologist workstations.
4.6.5 | The basic imaging manipulation interface of the imaging solution shall be capable of measuring a COBB angle on the displayed image.
4.6.6 | Multiple COBB angle measurements on a single image shall be possible on all Radiologist workstations.
4.6.7 | The Radiologist PACS application shall be capable of measuring a region of interest (ROI) on the displayed image.
4.6.8 | The measurement shall reflect average, minimum,maximum and standard deviation of the density within the ROI.
4.6.9 | Multiple ROI measurements on a single image shall be possible on all Radiologist workstations.
4.6.10 | Shapes shall be available for the ROI measurement on each Radiologist workstation. Circular, Rectangular, freehand drawn with mouse
4.6.11 | The reference points for any measurements shall be editable by the user. It shall be possible for the user to make text annotations
4.6.12 | It shall be possible to delete measurements.
4.6.13 | Any image to which text or measurements have been added will automatically be stored as a key image.
4.7 | Quality Assurance
4.7.1 | Shall conform with the latest Mammography Quality Standards Act (MQSA) regulations
4.7.2 | Quality assurance calibration shall be built into monitors to perform tasks such as luminance measurement and other quality assurance tests that are needed from radiation control for diagnostic monitors. If calibration tools are not build into monitors then calibration tools shall be included with diagnostic monitors.
4.7.3 | AAMP TG 18 test patterns to be included with PACS and shall always be available on cache storage, it shall also be able to view dually on both diagnostic monitors.
4.7.4 | L-angle distance patterns shall be included with PACS and shall aslo be available on on-line storage.
4.8 | Worklist Management
4.8.1 | The Radiologist PACS application shall support configurable work list management.
4.8.2 | The Radiologist PACS application shall have the contingency measure to search the PACS database for patient's Images, outside of the RIS application.
4.8.3 | The PACS application shall allow the user to search on the following DICOM fields from within the PACS application:
4.8.4 | Patients Last Name
4.8.5 | Patient Unique ID number or (MRN)
4.8.6 | Study Date
4.8.7 | Accession Number
4.8.8 | Study Description
4.8.9 | Referring Physician
4.8.10 | Modality
4.9 | The Radiologist PACS application shall allow for pattern searching within each of these fields. (Pattern searching refers to searching for a specific pattern of text or numerals, or symbols within the field).
4.9.1 | The Radiologist PACS application shall allow the User to use wildcards within the search criteria.
4.9.2 | The presentation of the work list to the user shall be configurable.
4.9.3 | The columns that are presented to the user should be configurable based on the DICOM tag required.
4.9.4 | The columns that are presented to the user shall be able, to be sorted ascending or descending.
4.9.5 | The order in which the columns are displayed should be configurable. Please give details.
4.9.6 | The Radiologist PACS application shall open the appropriate patient images on the viewing monitors after the patient has been selected from the work list.
4.9.7 | The Radiologist PACS application work list shall display the overall Patient examination history once a patient has been selected.
4.10 | Hanging Protocols
4.10.1 | When using the basic image manipulation app on a multi-monitor workstation, the user will be able to use hanging protocols. When using a single-monitor display it will be displayed according to the last and highest ranking presentation state, saved by a previous user.
4.10.2 | The Hanging Protocols shall be configurable.
4.10.3 | User defined hanging protocols shall be available to A specific users
4.10.4 | System wide define hanging protocols shall be available to all users.
4.10.5 | The configurations of the Hanging protocols shall include the ability to change the screen layouts to user define layouts.
4.10.6 | It shall be possible to change the layouts per monitor.
4.10.7 | It shall be possible to change the layouts across monitors. Hanging protocols must load selected image display features automatically, for instance-reference lines, synchronisation, zoom and pan,mirror image.
4.10.8 | It shall be possible to configure what is displayed in each segment of the display layout selected. Hanging protocols must load the advanced image manipulation application, and it must be a multi-modality hanging protocol.
4.10.9 | Which images are hung in which segment shall be configurable based on the following criteria.
4.10.10 | Study date
4.10.11 | Study time
4.10.12 | Body part
4.10.13 | The presence of previous patient
4.10.14 | examinations
4.10.15 | Modality
4.10.16 | Windows level
4.10.17 | Examination type
4.10.18 | Presentation states
4.10.19 | Prior vs curent
4.10.20 | Series description
4.10.21 | Study description
4.10.22 | Body position
4.10.23 | Laterality
4.10.24 | The Radiologist PACS application shall allow the User to save and label presentation states.
4.10.25 | These presentation states shall be saved as a presentation state object as part of the DICOM record in the PACS system
4.10.26 | The user shall be able to change the font size and colour of the annotations.
4.10.27 | The user shall be able to add "arrow" annotations to the image.
4.10.28 | Multiple annotations shall be possible on a single image.
4.10.29 | A pre defined list of annotations shall be available for the user.
4.10.30 | The user shall be able to edit the positions of an annotation.
4.10.31 | It shall be possible to show/hide annotations
4.10.32 | Spine annotations shall be possible in which the annotation automatically changes based on the previous annotation .e.g. L1 automatically changes to L2.
4.10.33 | Please state any other specific annotation functions your software solution provides.
4.10.34 | Annotations shall be saved as an overlay on the image.
4.11 | DICOM PRINTING
4.11.1 | Authorized users may print to a DICOM printer from the basic image manipulation application.
4.11.2 | The Radiologist PACS application DICOM print interface shall allow the user to select customizable layout formats.
4.11.3 | Please list the available print layouts your PACS application supports.
4.11.4 | The Radiologist PACS application DICOM print interface shall allow the user to Select different film sizes.
4.12 | Key Images
4.12.1 | Any images with annotations and measurements will automatically be saved as key images.
4.12.2 | These key images shall be saved as key images objects as part of the DICOM record in the PACS system.
4.13 | DESKTOP INTEGRATION BETWEEN RIS AND PACS (Institution Depended)
4.13.1 | The Radiologist PACS applications shall integrate seamlessly with the RIS application. A RIS/PACS solution is required from the vendor with seamless integration to HIS. If HIS is not avaialbale at institution, provision will be made for integration once HIS is deployed. i.e the institution shall be future proof for HIS integration.
4.13.2 | This integration shall automatically display the correct patient images in the PACS application when the patient is selected in the RIS.
4.14 | MULTIPLANAR RECONSTRUCTION (MPR)
4.14.1 | It shall be possible to construct a MPR view of CT or MRI studies from within the Radiologists PACS software.
4.14.2 | The user shall be able to see at least three frames
4.14.3 | The three frames shall be populated by a sagittal, coronal, and axial view of the loaded study.
4.14.4 | Positional correction – Each plane is defined by its relative position in two other planes. E.g. an axial plane is defined by its relative position in both the coronal (to define the left to right angulations) and the sagittal plane (to define the anterior to posterior angulation).
4.14.5 | The PACS application shall show two orthogonal lines per plane, demonstrating the relative position of the other two planes on the current plane i.e. reference lines.
4.14.6 | The user shall be able to view the orthogonal lines on all planes at all times.
4.14.7 | The user shall be able to hide/show the orthogonal lines.
4.14.8 | The user shall have an indication of which Orthogonal line is referencing which plain. The angle between the plains shall always be 90 Degrees in MPR mode.
4.14.9 | It shall be possible for the user to change the Position of the orthogonal lines dynamically.
4.14.10 | When an orthogonal line is changed the corresponding plane shall update in real time.
4.14.11 | It shall be possible to create a slab of data in any plane.
4.14.12 | The thickness of the volume shall be able to be dynamically adjusted by the user.
4.14.13 | The user shall be able to dynamically adjust the thickness of the volume.
4.14.14 | The user shall be able to choose between the following different intensity projections of the data within the volume.
4.14.15 | MIP (Maximum Intensity Projection) shall be possible.
4.14.16 | AveIP (Average Intensity Projection) shall be possible.
4.14.17 | MinIP (Minimum Intensity Projection) shall be possible
4.15 | Other Multiplanar Reconstruction Modes
4.15.1 | Multiplanar Reconstruction (MPR) Double Oblique
4.15.2 | It shall be possible for the user to perform MPR Double Oblique function within the Radiologist PACS Application.
4.15.3 | The only difference between the MPR Double Oblique function and the standard MPR function shall be that the angle between the planes can change.
4.15.4 | All other function describe in the standard MPR requirements shall be possible in the Double Oblique modes as well.
4.15.5 | Please state if any functions are not possible.
14.16 | Multiplanar Reconstruction (MPR) Curved Oblique
14.16.1 | It shall be possible for the user to preform a MPR Curved Oblique function within the Radiologist PACS Application.
14.16.2 | In the curved MPR Mode the user shall be able to see at least four frames.
14.16.3 | The frames shall contain the sagital, coronal, and axial planes. This will be used to place the points to create the curved oblique plane.
14.16.4 | The fourth frame shall contain the curved oblique plane.
14.16.5 | The user shall be able to mark points on any of the planes.
14.16.6 | The line connecting the points shall be created automatically.
14.16.7 | The user shall be able to edit the position of any of the points already displayed.
14.16.8 | The user shall be able to rotate the curved MPR using the curved oblique line as the axis of rotation
14.16.9 | It shall be possible to save any MPR from any mode to the PACS application for future review.
14.17 | VOICE RECOGNITION
14.17.1 | The voice recognition shall seamlessly integrate with the PACS and RIS.
14.17.2 | The voice recognition shall be set up for each individual user/radiologist to enable the system to learn from each unique speech of the users/radiologists.
14.17.3 | When a radiologist is no longer present at the healthcare facility, it shall be possible to revoke his speech license and make it available for a new user.
14.17.4 | There shall be one floating license that can be used by more than one user.
14.17.5 | The solution shall be able to convert Conversational speech into the highest accurate written text on a PACS report.
14.17.6 | There shall be no need for restrictive voice commands for punctuation or formatting.
14.17.7 | The solution shall have an Automatic and continuous learning system.
14.17.8 | The voice recognition shall be able to adjust to a variety of indigenous accents to ensure best results.
5 | CD/DVD Robot and workstation
5.1 | The following specification is noted for the CD/DVD robot solution set. This CD/DVD robot solution will be used to produce labelled DICOM CDs for external distribution. The specification includes the CD/DVD robot, workstation, monitor, keyboard, mouse, ext. The specification also includes the software functionality and interoperability standards required for this device.
5.2 | The quantity of CD/DVD robot workstations shall depend on the requirements of the institution
5.3 | All individual modules required to provide the functionality as specified in this specification shall be included and licensed to the purchaser.
5.4 | The robot shall be capable of producing dual media i.e. both CDs and DVDs.
5.5 | The robot shall have a capability of storing blank disks for each media type.
5.6 | The robot shall be capable of printing a pre-configured label onto the different types of media.
5.7 | The CD/DVD robot solution shall support the IHE Portable Data for Imaging Integration Profile (PDI) as the actor - Portable Media Creator. Please provide proof of your IHE integration.
5.8 | When media is created by the CD/DVD robot a DICOM viewing application shall be automatically written onto the media.
5.9 | Specify the load the robot can carry- it must be at least 100 CD's per day
5.10 | Indicate the speed of burning, and if it has 2 burners in parallel.
5.11 | It shall be possible to track the progress of writing and printing the DICOM media in the GUI of the CD/DVD robot solution.
5.12 | It shall be possible to gain access to a list of previously written studies in the GUI of the CD/DVD robot solution.
6. | RIS Functional Requirements
6.1. | User access control, security and audit tracking
6.1.1 | The access to the RIS solution shall be controlled by user access control, security and user tracking. These functions shall be based on sound IT principles. The following section outlines the requirements for RIS access, security and tracking.
6.1.2 | The RIS solution shall restrict access to only authorised users.The RIS solution shall support individual user login with password authentication.
6.1.3 | The RIS solution shall support role-based access controls to restrict access to system functionality based on predefined user group privileges.
6.1.4 | The RIS solution shall provide an audit trail of all user actions performed within the RIS solution, these audit trails shall be accessible to authorised users.
6.1.5 | The RIS solution shall provide an audit trail of all user access to and modification of patient data, these audit trails shall be accessible to authorised users.
6.1.6 | The robot shall be capable of printing a pre-configured label onto the different types of media.
6.1.7 | The CD/DVD robot solution shall support the IHE Portable Data for Imaging Integration Profile (PDI) as the actor - Portable Media Creator. Please provide proof of your IHE integration statement
6.1.8 | It shall be possible to mark individual patients records as confidential and restrict access to a specific user group.
6.1.9 | If the HIS is down and/or offline and the authorised user can't find the patient's record in the RIS, it shall be possible to register basic demographics for the patient into the RIS.
6.1.10 | It shall be possible for an authorised system administrator to define which fields to display in the registration process.
6.1.11 | It shall be possible to add a temporary MPI identifier into the master patient index identifier field in this case.
6.1.12 | It shall be possible to scan external documents and attached them to the patient record.
6.1.13 | It shall be possible to upload external documents and attached them to the patient record.
6.1.14 | It shall be possible to store multiple different types of files, such as Word, Jpeg, Tiff, PDF, ext. in the RIS.
6.2 | Basic order entry
6.20.1 | It is noted that traditional order entry will take place in the Hospital Information System (HIS) the functional requirements, in this case, are catering for the following scenarios. 1) In the case, the HIS is down and offline and an order is required to perform the procedure and/or procedure steps. 2) An additional procedure is required to be ordered by an authorised user such as Radiologists or Radiographer.
6.20.2 | It shall be possible to receive HL7 ORM type messages from the HIS and process these messages into the RIS.
6.20.3 | It shall be possible for an authorised user to manually order a radiology procedure request, in the RIS.
6.20.4 | It shall be possible for authorised users to search for existing patients already registered in the RIS database.
6.20.5 | It shall be possible to search on the following fields. Patient Master Index (PMI) identifier, Patient alternative identifier (state ID), Patient name, Patient DOB, Cell Phone number. Please give a list of the fields available to search.
6.20.6 | The search results shall display the following fields to assist in confirm patient identity. Patient Master Index (PMI) identifier, a Patient alternative identifier (state ID), Patient name, Patient DOB, Cell Phone number, Address.
6.20.7 | It shall be possible for an authorised user to select the requested procedure via a pre-defined procedure list. The pre-defined procedure lists shall support singular and multiple procedures steps.
6.20.8 | The pre-defined procedure lists shall support the South African government UPFS coding system.
6.20.9 | The pre-defined procedure shall be searchable on procedure code, modality type and procedure description or a combination of them.
6.20.10 | It shall be possible for an authorised user to add the priority of the order.
6.20.11 | It shall be possible for an authorised user to add clinical relevant data to the order via pre-defined drop-down lists, radio buttons and plain text fields.
6.20.12 | These data entry fields shall be configurable. Please give details of the types of fields and mechanisms available to configure the order entry clinical data components.
6.20.13 | It shall be possible to select the referring clinician who requested the exam from a pre-configured list. The list should include generic referring clinician information like name, contact details, email address, ext.
6.20.14 | In the case that the referring clinician's name is not on the pre-configured list it shall be possible for the user to add the referring clinician details to the RIS.
6.20.15 | It shall be possible for authorised users to scan and upload appropriate documentation, like procedure request forms, and attach these to the order.
6.20.16 | It shall be possible to add a patient's ambulatory status during the order entry process. i.e. bed, wheelchair, walking, ext.
6.20.17 | It shall be possible to mark individual patients records as confidential and restrict access to a specific user group.
6.3 | Procedure Validation
6.3.1 | Procedure validation is the step associated with validating, that the exam that has been ordered is appropriate, based on the patient clinical history, previous examinations, clinical contraindications, reason for exam, ext.
6.3.2 | It shall be possible for an authorised user to validate if the procedure is appropriately ordered.
6.3.3 | It shall be possible to pre-configure which procedures need to be validated.
6.3.4 | The orders containing procedures which require validation should appear on a validation worklist.
6.3.5 | The following information shall be available to the authorised user who is performing the procedure validation:
6.3.6 | Access to patients previous imaging and reports.
6.3.7 | Integrated link to call the images and display images via the PACS solution.
6.3.8 | Referring clinicians scanned request form when the order was manually entered.
6.3.9 | Referring clinicians clinically relevant data if the order was placed electronically and sent from the HIS
6.3.10 | Other scanned documents which are appropriate to the patient and/or the ordered procedure.
6.3.11 | It shall be possible for the authorised user to reject or approve the procedure.
6.3.12 | In the case of rejecting the procedure, it shall be possible for the authorised user to add a note with the reason for rejecting the procedure.
6.3.13 | It shall be possible for the authorised user to change the procedure ordered.
6.3.14 | It shall be possible for the authorised user to add a procedure to the order.
6.3.15 | It shall be possible for the authorised user to change the priority of the order.
6.3.16 | An email or SMS notification shall be sent to the referring clinician, who ordered the procedure, once the procedure is validated. The email should inform the referring clinician of the outcome of the validation process. i.e. rejected or approved.
6.4 | Order Scheduling
6.4.1 | Scheduling is the workflow step associated with scheduling an exam for a future date. In the South African public sector patients do not adhere to time-based scheduling due to various constraints like transport. So the solution needs to accommodate for slot based scheduling. Slot-based scheduling is based on having a number of slots allocated per day rather than specific times during the day. Order scheduling must be as follows:
6.4.2 | The RIS shall allow an authorised user to schedule orders for a future appointment date.
6.4.3 | Slot-based scheduling shall be preferred to time-based scheduling.
6.4.4 | It shall be possible to pre-configure which procedures require scheduling.
6.4.5 | Orders that include procedures which require scheduling, shall appear on a scheduling worklist.
6.4.6 | It shall be possible to filter the worklist by modality.
6.4.7 | It shall be possible to filter the worklist by the department, for example, emergency department, or orthopaedic department, ext.
6.4.8 | It shall be possible to filter or sort the worklist by priority.
6.4.9 | Multiple procedure steps under one requested procedure shall remain grouped together during scheduling.
6.4.10 | It shall be possible to pre-configure time slots per modality, per day.
6.4.11 | It shall be possible to allocate a pre-configured number of slots used per procedure step.
6.4.12 | It shall be possible to configure the number of available slots per modality per session. A session is a period of time for example 8h00 to 12h00 (morning session) 12h00 to 16h00 (afternoon session).
6.4.13 | It shall be possible to block out slots making them unavailable for scheduling.
6.4.14 | It shall be possible to create pre-configured procedure notes per procedure which give additional details about the procedure.
6.4.15 | The solution shall automatically calculate which is the best slot to schedule the procedure and procedure step, based on the modality, procedure type, and slot availability.
6.4.16 | The solution shall allow authorised users to manually schedule, procedures and procedure steps.
6.4.17 | It shall be possible for an authorised user to double book procedures and procedure steps.
6.4.18 | The authorised user shall have access to an overview of the scheduling calendars per modality or multiple selected modalities.
6.4.19 | It shall be possible for an authorised user to find a patient's scheduled procedure and/or scheduled procedure step and reschedule it.
6.4.20 | It shall be possible to add a note why the procedure or procedure step has been rescheduled.
6.4.21 | It shall be possible for an authorised user to find a patient's scheduled procedure or procedure step to cancel it.
6.4.22 | It shall be possible for an authorised user to record the reason for the cancellation.
6.4.23 | It shall be possible for an authorised user to print, email or SMS, scheduled details to the patient and/or the referring clinician once the procedure is scheduled.
6.4.24 | It shall be possible for an authorised user to print daily scheduled lists.
6.5 | Patient Waiting Worklist
6.5.1 | The patients waiting worklist identifies patients that have been scheduled for that day and are expected to arrive for their procedure. There are a number of different scenarios of how a patient may arrive at the department. The patient may be a walk in an (unscheduled) patient, scheduled patient, emergency patient, ward patient, or theatre patient. The specifications for the features relating to the patient arrival worklist are detailed below:
6.5.2 | It shall be possible for an authorised user to view the patients waiting worklist.
6.5.3 | All orders which have been ordered electronically or manually and don't require validation or scheduling shall appear on a waiting worklist.
6.5.4 | Orders which have been scheduled for that day shall appear on the patients waiting worklist.
6.5.5 | It shall be possible to filter the worklist by modality.
6.5.6 | It shall be possible to filter the worklist by the department, for example, emergency department, or orthopaedic department, ext.
6.5.7 | It shall be possible to filter or sort the worklist by priority.
6.5.8 | It shall be possible for an authorised user to rescheduled a procedure from the patients waiting worklist.
6.5.9 | It shall be possible to record the reason why the procedure has been rescheduled.
6.5.10 | It shall be possible for an authorised user to cancel a procedure from the patients waiting worklist.
6.5.11 | It shall be possible to record the reason why the procedure has been cancelled.
6.5.12 | It shall be possible for an authorised user to indicate that the patient has arrived and is waiting for the exam to be performed.
6.6 | Performed Procedure Worklist
6.6.1 | The performed procedure worklist identifies when the patient has arrived in the department and is waiting for the procedure to be performed. This is essentially the radiographer worklist. The specifications for features relating to the procedure performed workflow step are detailed below:
6.6.2 | It shall be possible for an authorised user to view the performed procedure worklist.
6.6.3 | All procedure steps of the procedure to be performed shall be available on a DICOM modality worklist and available for the modality to query.
6.6.4 | It shall be possible to filter the worklist by modality.
6.6.5 | It shall be possible to filter the worklist by the department, for example, emergency department, or orthopaedic department, ext.
6.6.6 | It shall be possible to filter or sort the worklist by priority.
6.6.7 | Once a user has indicated that the procedure has started the RIS shall automatically assign that procedure to the performing user.
6.6.8 | Once a procedure or procedure step has been assigned to a user no other user shall be able to assign that procedure.
6.6.9 | Assigned procedures shall be transferable between authorised users.
6.6.10 | The RIS shall display patient related alerts and warning such as allergies, ext.
6.6.11 | The RIS shall warn the user if the patient is scheduled for multiple procedures for the same day across modalities.
6.6.12 | The RIS shall have the functionality to capture procedure related questioners.
6.6.13 | The RIS shall allow specific procedure questioners to be mandatory. i.e. the procedure can't be started or completed until the questionnaire is completed by the authorised user.
6.6.14 | It shall be possible for an authorised user to cancel a procedure from the performed procedure worklist.
6.6.15 | It shall be possible for an authorised user to add a note containing the reason for the cancellation, to the system.
6.6.16 | It shall be possible for an authorised user to rescheduled a procedure from the performed procedure worklist.
6.6.17 | It shall be possible for an authorised user to add a note containing the reason for the rescheduling, to the system.
6.6.18 | It shall be possible for an authorised user to order an additional procedure from the performed procedure worklist.
6.6.19 | It shall be possible for an authorised user to update the originally ordered procedure to indicate the actual procedure performed.
6.6.20 | It shall be possible to add multiple responsible users details (i.e. Radiographers names) to the performed order.
6.6.21 | It shall be possible for authorised users to access the patient's imaging history including previous images and reports via an integrated link to the PACS.
6.6.22 | The integration shall be seamless with no need for the authorised user to retype username or password to gain access to the PACS.
6.6.23 | It shall be possible for an authorised user to view scanned documents and ordering clinician request forms.
6.6.24 | It shall be possible for an authorised user to add clinical notes to the performed procedure.
6.6.25 | It shall be possible for an authorised user to mark the procedure as completed.
6.7 | Reporting Worklist
6.7.1 | The reporting worklist identifies when the procedures and procedure steps are completed and the procedure is ready for reporting. Essentially this is the Radiologists reporting worklist. It should be noted that the solution will be deployed into a teaching environment, and such shall cater to the registrar and consultant reporting workflows. The specifications for features relating to the reporting workflow step are detailed below.
6.7.2 | All procedures and procedure steps which have been marked completed by the Radiographer and not flagged as "not for reporting" shall appear on a reporting worklist.
6.7.3 | It shall be possible for an authorised user to filter the reporting worklist by modality.
6.7.4 | It shall be possible to filter the worklist by the department, for example, emergency department, or orthopaedic department, ext.
6.7.5 | It shall be possible to filter or sort the worklist by priority.
6.7.6 | The solution shall display patient related alerts and warning such as allergies, ext.
6.7.7 | The solution shall highlight if the patient is scheduled for multiple procedures for that day across modalities.
6.7.8 | On opening an order to view images and create a report the solution shall lock that order so no other user can report on it.
6.7.9 | Users shall have access to orders opened by another user but in a read-only mode and the user shall be warned that the order is already opened by another user.
6.7.10 | On opening an order to start reporting the linked images in the PACS system shall automatically be called and displayed on the 2 diagnostic monitors.
6.7.11 | The integration shall be seamless with no need for the authorised user to retype username or password to gain access to the PACS.
6.7.12 | The authorised user shall have access to all the clinical data relevant to the order, this includes, patient age, priority, referring clinician request, clinical history, the reason for exam, scanned documents, radiographer notes. Ext.
6.7.13 | The solution shall support a voice recognition workflow where dictations are automatically transcribed to text and a registrar or consultant can edit their own reports.
6.7.14 | The solution shall support pre-configured report templates. These shall be configurable on a global or user level.
6.7.15 | It shall be possible to assign templates to procedures or procedure steps.
6.7.16 | It shall be possible to filter templates by modality.
6.7.17 | It shall be possible for authorised users to call report templates via voice commands. i.e. saying "CT normal chest" will import the CT normal chest pre-configured template.
6.7.18 | The solution shall support stop points within the pre-configured report templates. Stop point are pre-configured points in a template where registrars or consultants can add measurements or comments and then move quickly to the next stop point in a template.
6.7.19 | It shall be possible for an authorised user to add an ICD10 diagnosis code to the procedure.
6.7.20 | The solution shall support pre-loaded ICD10 codes based on South African accepted ICD10 code lists.
6.7.21 | Possible for an authorised user to add BI-RADS codes to the procedure in case of mammography.
6.7.22 | It shall be possible for a registrar to sign off their report after self-editing.
6.7.23 | A registrars report shall be signed off as a preliminary report. This shall be clearly indicated on the report and the reporting registrars name shall appear on the report.
6.7.24 | A registrar signing off their report shall trigger an HL7 message to send the preliminary report to HIS and PACS.
6.7.25 | It shall be possible for a consultant to sign off their report after self-editing.
6.7.26 | A consultant's report shall be signed off as a final report. This shall be clearly indicated on the report and the consultant's name shall appear on the report.
6.7.27 | A consultant signing off their report shall trigger an HL7 message to send the final report to HIS and PACS.
6.7.28 | It shall be possible for authorised consultant users to edit registrars preliminary report and sign them off as final reports. In this case, the final report shall be indicated as such and have both the registrars and consultants name on the report.
6.7.29 | Once a report has reached final status it can't be changed or edited. If changes are required, It shall be possible to add an addendum to the report.
6.7.30 | The addendum shall appear on top of the original report.
6.7.31 | It shall clearly indicate that the report is an addendum.
6.7.32 | The solution shall allow authorised users to open a second study (e.g. STAT Exam) while another study is open, without closing or losing any work done on the originally opened study?
6.8 | Report distribution
6.8.1 | The solution shall support the distribution of the report via the following methods.
6.8.2 | Web based application browser viewer
6.8.3 | Secure Email
6.8.4 | Secure access through clinicians portal
6.8.5 | Printed report
6.8.6 | HL7 outbound messaging to HIS
6.9 | Quick Patient Registration
6.9.1 | The process of registering a patient quickly in the RIS in the case the HIS is down and/or an examination needs to be done in an emergency.
6.9.2 | If the HIS is down and/or offline and the authorised user can't find the patient's record in the RIS, it shall be possible to register basic demographics for the patient into the RIS. It shall be posible to backdate registration information
6.9.3 | This registration shall be based on a quick registration process with a set of minimum required fields.
6.9.4 | It shall be possible for an authorised system administrator to define which fields to display in the registration process.
6.9.5 | It shall be possible for an authorised system administrator to define which fields in the registration process are mandatory.
6.9.6 | It shall be possible to add a temporary MPI identifier into the master patient index identifier field in this case.
6.9.7 | It shall be possible to scan external documents and attached them to the patient record.
6.9.8 | It shall be possible to print patient ordering slips
6.9.9 | It shall be possible to upload external documents and attached them to the patient record.
6.9.10 | It shall be possible to printing demographic labels
6.9.11 | It shall be possible to store multiple different types of files, such as Word, Jpeg, Tiff, PDF, ext. in the RIS. Please give a list of file types supported by your RIS.
7. | PACS Functional Requirements
7.1 | User access control, security and audit tracking
7.1.2 | The access to the PACS solution shall be controlled by user access control, security and user tracking. These functions shall be based on sound IT principles. The following section outlines the requirements for RIS access, security and tracking:
7.1.3 | The PACS solution shall restrict access to only authorised users.
7.1.4 | The PACS solution shall support individual user login with password authentication.
7.1.5 | The PACS solution shall support role-based access controls to restrict access to system functionality based on predefined user group privileges.
7.1.6 | The PACS solution shall provide an audit trail of all user actions performed within the PACS solution, these audit trails shall be accessible to authorised PACS users.
7.1.7 | The PACS solution shall provide an audit trail of all user access to and modification of patient data, these audit trails shall be accessible to authorised PACS users.
7.1.8 | It shall be possible to mark individual patients records as confidential and restrict access to a specific user group.
7.2 | Data Management
7.2.1 | Data management is the functionality in the PACS which allows authorised user to manage the data lifecycle of data stored in the archive. This includes moving data from expensive storage hardware to less expensive storage hardware based on a set of data lifecycle rules. It also includes managing retention periods of studies and when and which studies can be deleted.
7.2.2 | The PACS shall support rule-based lifecycle data management, to manage the movement of data in a tiered archive solution.
7.2.3 | The PACS shall support rule-based lifecycle data management, to manage the retention of data in the archive solution. Retention of data is achieved by not allowing users the function to delete studies if the study is protected by the retention rule.
7.2.4 | The PACS solution shall support automatic rule-based routeing of data to different nodes.
7.2.5 | The basic PACS functionality is functionality that is traditionally provided by the PACS solution and applicable to all authorised users. The specifications for base PACS functionality are detailed below.
7.2.6 | It shall be possible to change window level and centre.
7.2.7 | It shall be possible to store pre-set window settings for CT viewing.
7.2.8 | It shall be possible to change the image layout on the viewing screen.
7.2.9 | It shall be possible to measure distance, angle, Cobb angle, a Region of interest, Hounsfield units, ext.
7.2.10 | It shall be possible to manually scroll through large cross-sectional studies like CT, Angiography and MRI.
7.2.11 | It shall be possible to automatically scroll through large cross-sectional studies with a cine loop function.
7.2.12 | It shall be possible to vary the speed and direction of the cine loop.
7.2.13 | It shall be possible to link multiple series in a cross-sectional study like CT and MRI and scroll through the series simultaneously.
7.2.14 | It shall be possible for an authorised user to create custom shortcut keys for commonly used functions.
7.2.15 | It shall be possible for an authorised user to export images as jpeg images.
7.2.16 | It shall be possible for an authorised user to export images as Tiff images.
7.3 | Radiographers PACS Functionality (INSTITUTION OPTIONAL)
7.3.1 | The Radiographer PACS functionality refers to PACS functionality that is required, which is specifically related to Radiographers. This functionality should be available to authorised radiography users working on any workstations in the department, especially the RIS workstations which are located next to or in the region of modalities. The specifications for PACS features relating to radiographers are detailed below.
7.3.2 | The PACS shall provide for pre-configured text labelling
7.3.3 | The PACS shall allow authorised users to label and save presentation states.
7.3.4 | The PACS shall allow authorised users to print DICOM films to the DICOM laser printer.
7.3.5 | It shall be possible for an authorised user to send DICOM studies to the CD/DVD robot. (Institution Optional)
7.3.6 | The Radiographers PACS application shall allow The PACS Administrator to correct mismatched study information between the RIS application and the PACS application.
7.3.7 | The PACS Administrator PACS applications shall allow the user the ability to insure database integrity between the data on the RIS and data on the PACS system.
7.4 | Radiologists PACS Functionality
7.4.1 | The Radiologists PACS functionality refers to functionality that is required to view and diagnose radiology images in the radiology department. This functionality should be available to radiologists working on dedicated diagnostic workstations in the department and at home when they are on call. The specifications for PACS features relating to radiologists are detailed below.
7.4.2 | The PACS shall allow authorised users to create user-specific function keys for commonly used functions Please give details of which functions are pre-set and which are configurable if any.
7.4.3 | The PACS shall allow authorised users to select and store key images.
7.4.4 | Key images stored in the PACS shall be stored as DICOM key image objects.
7.4.5 | Annotations, presentation overlays, presentation states stored as part of the DICOM study shall be visible to the authorised user.
7.4.6 | The PACS shall allow authorised users to add annotations, label and save presentation states.
7.4.7 | Presentation states saved to the PACS shall be stored as DICOM presentation state objects.
7.4.8 | The PACS shall allow authorised users to configure user-specific hanging protocols. Please give details of the mechanisms used to deliver this functionality.
7.4.9 | The solution shall provide the authorised users with a list of recently viewed studies.
7.4.10 | It shall be possible for authorised users to flag interesting cases and save them in a teaching file folder.
7.4.11 | It shall be possible for authorised users to store interesting cases.
7.4.12 | The solution shall automatically link 2, 3 or 4 series within a cross-sectional type study based on a predefined hanging protocol and synchronise scroll through the images.
7.4.13 | The solution shall allow an authorised user to manually link 2, 3 or 4 series within a cross-sectional type study and synchronise scroll through the images.
7.4.14 | The solution shall provide for remote reporting services by allowing a user to login to the system and access functionality remotely. This will include creating reports and accessing images.
7.5 | PACS Advanced Functionality
7.5.1 | The PACS solution shall provide basic advanced visualization tool applicable to volumetric studies as specified below:
7.5.2 | The PACS functionality shall support Multi-Planar reconstructions (MPR) in the following planes, axial, coronal, sagittal, and oblique.
7.5.3 | The PACS functionality shall support Curved-Planar reconstructions (CPR) including projected, straightened and stretched views.
7.5.4 | The PACS functionality shall support changing slab slice thickness.
7.5.5 | The PACS functionality shall support slab rendering based on maximum (MIP), minimum (MinIP) and average (AvgIP) intensity projections.
7.5.6 | The PACS functionality shall support 3D volume rendering for specific anatomy parts such as lung, bone, vessels and skin.
7.5.7 | The PACS functionality shall support rendering 3D animations and exporting these animations as AVI files.
7.5.8 | The PACS functionality shall support basic segmentation tools, including bone removal, cropping, clipping, freehand selection.
7.5.9 | The PACS functionality shall support using selected advanced visualization tools automatically in predefined hanging protocols.
7.5.10 | The PACS functionality shall support volume measurements.
7.5.11 | The PACS shall support basic dental specific functionality, please give a detailed description of your solutions specific dental functionality. (Institution Specific)
7.5.12 | The PACS must have compatability for possible future Artificial Intelligence (AI) applications.
8. | Clinicians PACS Functionality
8.1 | One of the critical functions the solution shall support is a clinician portal. This is the portal that the clinicians will use to access order scheduling information, images and reports. The DoH is looking to take advantage of the new mobile technologies and zero footprint viewers throuhg a web based clinicians portal
8.2 | The solution shall support viewing of order information, images and reports through a web-based clinicians portal.
8.3 | It shall be possible for authorised users to search for patients based on their Patient Master Index identifier, Patient name, study date, and/or study date range
8.4 | In the case of reports created by registrars and signed off by a consultant, both the preliminary and final report shall be available.
8.5 | The solution shall support basic PACS functionality as noted through the clinicians portal.
8.6 | The solution shall support basic visualisation tools like MPR, MIP and 3D rendering through the clinicians portal.
8.7 | The solution shall support access to the portal via mobile technology such as smartphones, and tablets or on a browser through URL. Please give details of what browsers are compatiable.
8.8 | The solution shall support cross-browser viewing.
8.9 | The solution shall support Android devices.
8.10 | The solution shall support Apple devices.
8.11 | The solution shall support Windows devices.
8.12 | The solution shall support access to images through an HTML5 web browser.
8.13 | The data streamed from the PACS or RIS archive to the device shall be based on a secure encrypted service.
8.14 | No image or report data shall remain stored on the device once the connection is dropped or disconnected.
8.15 | For the transcriptionist provision shall be made for:
8.16 | Transcriptionist Workstations (Institution optional)
8.17 | USB transcriptionist set including
8.18 | headset
8.19 | and paddle (Institution optional)
8.20 | Provision shall be made based on an institutional requirements
8.21 | RIS Workstations
8.22 | PACS Administrator Workstation and
8.23 | or Notebook and or Tablet
8.24 | Orthopeadic Workstation
8.25 | Radiogrpaher Viewing Workstation
8.26 | All Network equipment shall be able to connect 1Gbs Network.
8.27 | All computer equipment shall include keyboard and mouse.
8.28 | Should the institution not have the necessary network requirements for the system to operate optimally, a network quotation will be supplied
8.29 | (Institution optional)
8.30 | The following should be considered when developing the network quotation:
8.31 | The network shall have a fiber uplink to the core switch located in the server room at a minimum of 1Gb/s to connect to FSHealth network.
8.32 | Layer 3 manageable switch is required with PoE with redundant/spare switch available.
8.33 | Current cabinets may be utilized and all newly installed cabling must be marked clearly.
8.34 | A site survey is recommended before the proposal is submitted to ensure the correct equipment is installed.
8.2 | The final proposal must be approved by End Users
8.2.1 | Shall provide a full diagram and description of their whole solution including RIS-PACS, integration with HIS (institution optional), data migration,archiving options, server details,primary Archive .The PACS shall be able to send images to other institutions and to a future central data repository or VNA, it shall also be able to receive images from other institutions i.e. the PACS shall include the DICOM Service Class license Query/Retrieve SCU and SCP.
8.2.2 | The solution offered shall make provision for 5 years online (cached) storage.
8.2.3 | In the event of a failure within the system, the system will automatically alert the PACS administrator and vendor of potential problems. A Total pro-active self-monitoring system of server room shall be provided and alerts shall also be done in the form of sms or any other messaging system to alert PACS administrator and vendor. Monitoring shall form part of contract.
8.2.4 | A full Disaster Recovery solution shall be provided. This implies that a separate archive solution shall be provided in a separate location within the hospital grounds or elsewhere. In the case of a complete disaster of the primary archive the disaster recovery solution shall provide the radiology department with a working archive solution seamlessly. This solution shall provide the same functionality as the primary archive.
8.2.5 | Shall be responsible to liaise with modality suppliers and ensure correct connection of all modalities to the PACS server. All modalities shall make useof modality worklist.
8.2.6 | Shall ensure smooth transition from hospital power supply to UPS power when there is power-outages. Switch over configuration shall be seamless to switch over and uninterrupted. UPS for RIS/PACS servers to be responsibility of vendor. UPS shall be able to maintain servers for minimum an hour.
8.2.7 | Shall supply their own UPS’s - independent of the Departmental UPS’s currently available
8.2.8 | Shall supply 2 X aircons per server room for the main and secondary server rooms, where 1 aircon will be used and the other will be off asa redundant measure. Vendors shall maintain aircons at their own cost.
9 | Data Mining
9.1 | The solution shall provide access to data mining tools to pull statistical data from the system.
9.2 | It Must be possible to access workflow activities based on the nuber of procedures perfomed by the student radiographer, qualified radiographer, the number of repots generated by the registrar and or consultatnt the number of procedures done and the number of patients done.
9.3 | It must be possible to access data relating to waiting times and workflow steps.
9.4 | It must be possible to extract the list of procedures reported on by registrars name e.g registrars logbook.
9.5 | It must be possible to extract capacity and demand reports based on modality work areas's availability and usage in the scheduling system.
9.6 | All reports shall be able to be exported in CSV and or excel format with data types correctly formatted.
9.7 | All data mining capabilities offered in the solution must be descibed.
10. | Interoperability Requirements
10.1 | DICOM
10.1.1 | Provide DICOM conformance statement for the PACS solution offered.
10.1.2 | Provide DICOM conformance statement for the RIS solution offered. If the RIS is not providing any DICOM services, indicate this below.
10.1.3 | Provide the DICOM conformance statement for the CD/DVD Robot solution. (Institution Specific)
10.1.4 | The PACS solution shall support the DICOM storage service as SCP and SCU.
10.1.5 | The PACS solution shall support the DICOM query and retrieve service as SCP and SCU.
10.1.6 | The PACS solution shall support the DICOM print service as SCU.
10.1.7 | The PACS solution shall support the DICOM Modality Performed Procedure Step (MPPS) service as SCP.
10.1.8 | The PACS solution shall support the DICOM Storage commitment service as SCP. Provide evidence of required service by referencing the appropriate section in your DICOM Conformance Statement.
10.1.9 | The solution shall the DICOM Modality Worklist Service as SCP. This service may be implemented in the PACS or RIS solution.
10.1.10 | The solution shall support all required DICOM SOP classes to store DICOM objects based on the modalities listed herein.
10.2 | IHE
10.2.1 | The solution shall support the IHE Patient Demographics Query (PDQ) profile.
10.2.2 | The solution shall support the IHE Patient Identifier Cross-Referencing (PIX) profile.
10.2.3 | The solution shall support the IHE Consistent Time (CT) profile.
10.2.4 | The solution shall support the IHE Scheduled Workflow (SWF) profile.
10.2.5 | The solution shall support the IHE Patient Information Reconciliation (PIR) profile.
10.2.6 | The solution shall support the IHE Consistent Presentation of Images (CPI) profile.
10.2.7 | The solution shall support the IHE Audit Trail and Node Authentication (ATNA) profile with special reference to the Radiology defined options.
10.2.8 | Required to submit IHE integration statements for all IHE profiles which the solution supports. This should include IHE integration statements for the PACS, the RIS and the CD Robot solutions.
10.2.9 | The PACS shall adhere to the IHE profile: patient information reconciliation (PIR).
10.2.10 | The PACS shall adhere to the following IHE Standard: IHE Standard—Cross Community Access of IHE.
10.2.11 | The PACS shall adhere to the IHE profile: Cross-enterprise Document Sharing for Imaging (XDS-I.b).
10.2.12 | Bidders are required to submit IHE integration statements for all IHE profiles which the solution supports.
10.3 | HL7
10.3.1 | Bidders shall provide an original OEM HL7 integration document for the PACS solution which documents which HL7 message types, data types, and data fields are supported by your PACS solution.
10.3.2 | Bidders shall provide an original OEM HL7 integration document for the RIS solution which documents which HL7 message types, data types, and data fields are supported by your RIS solution.
10.3.3 | The solution shall support an inbound and outbound HL7 messaging interface. Provide evidence of requirement by referencing appropriate sections in your HL7 integration documents.
10.3.4 | The solution shall support the following HL7 message type
10.3.5 | ADT- Admission Discharge Transfer
10.3.6 | Provide evidence of requirement by referencing appropriate sections in your HL7 integration documents.
10.3.7 | DFT – Detailed Financial Transaction. Provide evidence of requirement by referencing appropriate sections in your HL7 integration documents.
10.3.8 | The solution shall support the following HL7 message type
10.3.9 | ORM- Order Management
10.3.10 | Provide evidence of requirement by referencing appropriate sections in your HL7 integration documents.
10.3.11 | The solution shall support the following HL7 message type
10.3.12 | ORR- Order Response (to any ORM)
10.3.13 | Provide evidence of requirement by referencing appropriate sections in your HL7 integration documents.
10.3.14 | The solution shall support the following HL7 message type
10.3.15 | ORU- Observation Report Unsolicited
10.3.16 | Provide evidence of requirement by referencing appropriate sections in your HL7 integration documents.
10.3.17 | Bidders are to provide a list of HL7 trigger points within the solution. A trigger point is a point at which the RIS or PACS will trigger an outbound HL7 message.
10.4 | Vendor Neutral Archieve (VNA)
10.4.1 | The institution may implemnt VNA
10.4.2 | PACS Solution shall support intergration into VNA solution in the future, bidders are to provide details of how their solution will maintain data intergrity, DICOM storage and synchronise radiology data to such a VNA.
11. | DATA MIGRATION
11.1 | It is the intention of the Hospital to migrate all data from the existing PACS& RIS solution to the new PACS & RIS solution. This includes Patient Demographics, Requests, Visits, Orders, Dictations, Reports, Users, Procedures Tables, Billing codes, Inventory, etc.
11.2 | Shall migrate the current RIS/PACS data to the new RIS/PACS Archive. It shall be possible to view the current PACS images in the new PACS archive and the same shall apply for the RIS records, it shall be possible to view current RIS records and reports on the new RIS/PACS solution. The migration process shall in no way effect clinical imaging/ workflow activities. The successful vendor shall be 100% complete with migration within 6 months
11.3 | Provide your data migration methodology statement.
11.4 | An integrity data check and test patients must be done by the new vendor.
11.5 | Corrupted data identified during the migration process, should be highlighted to the hospital and succesful bidder should provide a vaible solution to deal with corrupted data.
11.6 | Storage capacity at Hospital is as follows:
11.7 | -Storage in Archive:
11.8 | -Image PACS Redo:
11.9 | -Oracles Files:
11.10 | -Database Files:
11.11 | -Temp:
11.12 | All expenses of data migration including any dealings with 3rd parties are to be funded by succesful bidder.
11. | BUSINESS INTELLIGENCE
11.1 | PACS administrator shall be able to create/build reports from any location he/she has access to the PACS/RIS solution
11.2 | The business intelligence should provide online analytical processing (OLAP), statistical analysis, forecasting, and data mining.
11.3 | The business intelligence should provide management with real time reports and statistical analysis.
11.4 | Business intelligence shall be able to analyse data from modalities, RIS, PACS.
15.5 | Features and Functions OLAP (On-line analytical processing) Query and Reporting Advanced Analytics Collaborative BI Corporate Performance Management (Portals, Scorecards, Dashboards) Proactive BI (BAM) Clinical analytics